## All tables combined
EPA/625/R-96/010a
Compendium of Methods
for the Determination of
Inorganic Compounds
in Ambient Air
Compendium Method IO-4.2
DETERMINATION OF REACTIVE
ACIDIC AND BASIC GASES
AND STRONG ACIDITY OF
ATMOSPHERIC FINE PARTICLES
(
Center for Environmental Research Information
Office of Research and Development
U.S. Environmental Protection Agency
Cincinnati, OH 45268
June 1999
Method IO-4.2
Acknowledgments
This Method is a part of Compendium of Methods for the Determination of Inorganic Compounds in Ambient
Air (EPA/625/R-96/010a), which was prepared under Contract No. 68-C3-0315, WA No. 2-10, by Midwest
Research Institute (MRI), as a subcontractor to Eastern Research Group, Inc. (ERG), and under the
sponsorship of the U.S. Environmental Protection Agency (EPA). |  | Justice A. Manning, John O. Burckle,
Scott R. Hedges, Center for Environmental Research Information (CERI), and Frank F. McElroy, National
Exposure Research Laboratory (NERL), all in the EPA’s Office of Research and Development, were
responsible for overseeing the preparation of this method. |  | Other support was provided by the following
members of the Compendia Workgroup:
• | James L. Cheney, U.S. Army Corps of Engineers, Omaha, NE
• | Michael F. Davis, U.S. EPA, Region 7, KC, KS
• | Joseph B. Elkins Jr., U.S. EPA, OAQPS, RTP, NC
• | Robert D. Lewis, U.S. EPA, NERL, RTP, NC
• | Justice A. Manning, U.S. EPA, ORD, Cincinnati, OH
• | William A. McClenny, U.S. EPA, NERL, RTP, NC
• | Frank F. McElroy, U.S. EPA, NERL, RTP, NC
• | William T. "Jerry" Winberry, Jr., EnviroTech Solutions, Cary, NC
This Method is the result of the efforts of many individuals. Gratitude goes to each person involved in the
preparation and review of this methodology.
Author(s)
• | William T. "Jerry" Winberry, Jr., EnviroTech Solutions, Cary, NC
• | Thomas Ellestad, U.S. EPA, RTP, NC
• | Bob Stevens, U.S. EPA, RTP, NC
Peer Reviewers
• | Delbert Eatough, Brigham Young University, Provo, UT
• | Shere Stone, University Research Glassware Corp., Chapel Hill, NC
• | Petros Koutrakis, Harvard School of Public Health, Boston, MA
• | J. Waldman, Robert Wood Johnson Medical School, New Brunswick, NJ
• | Lauren M. Drees, U.S. EPA, NRMRL, Cincinnati, OH
DISCLAIMER
This Compendium has been subjected to the Agency's peer and administrative review, and it has
been approved for publication as an EPA document. Mention of trade names or commercial products
does not constitute endorsement or recommendation for use.
ii
Method IO-4.2
Determination of Reactive Acidic and Basic
Gases and Strong Acidity of Atmospheric
Fine Particles (
TABLE OF CONTENTS
Page
1. Scope . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  |  | 4.2-1
2. | Applicable Documents |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-2
| 2.1 | ASTM Standards . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-2
| 2.2 | Other Documents | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-2
3. | Summary of Method |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-2
4. | Significance | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-2
5. | Definitions . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-3
6. | Factors Affecting Denuder Efficiency |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-4
7. | Apparatus | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-4
| 7.1 | Sampling . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-5
| 7.2 | Analysis | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-7
8. | Reagents and Materials . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-8
9. | Preparation of Coating and Extraction Reagents |  |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . | 4.2-10
| 9.1 | Impactor Frit Coating Solution Preparation |  | . . . . . . . . . . . . . . . . . . . . . . . . . . | 4.2-10
| 9.2 | Impactor Frit Extraction Solution Preparation . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-10
| 9.3 | Annular Denuder Coating Solutions Preparation . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-10
10. | Elutriator and Acceleration Jet (Inlet) Assembly |  |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . | 4.2-10
11. | Impactor Frit Preparation and Installation |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-11
| 11.1 | Impactor Frit Installation | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-11
| 11.2 | Impactor Frit Preparation . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-11
12. | Filter Pack Preparation and Assembly . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-11
13. | Annular Denuder System Preparation |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-12
| 13.1 | Annular Denuder Coating Procedure . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-13
| 13.2 | Annular Denuder Drying Procedure . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-13
| 13.3 | Annular Denuder System (ADS) Assembly |  | . . . . . . . . . . . . . . . . . . . . . . . . . . | 4.2-13
| 13.4 | Laboratory Leak-Check of ADS . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-14
14. | Sampling . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-14
| 14.1 | Placement of Denuder System . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-14
| 14.2 | Start-Up | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-15
| 14.3 | Sample Shutdown | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-16
| 14.4 | Corrective Action for Leak Test Failure | . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-17
15. | ADS Disassembly |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-17
16. | Extraction Procedures |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-18
| 16.1 | Impactor Frit Coating Extraction (Optional) . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-18
| 16.2 | Denuder Extractions | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-18
| 16.3 | Filter Extraction . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-19
iii
TABLE OF CONTENTS (continued)
Page
17. | Ion Chromatography Analysis |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-20
| 17.1 | Standards Preparation | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-21
| 17.2 | Reagent Preparation . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-21
| 17.3 | Sample Preparation . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-22
| 17.4 | Basic System Operations - Start-up and Shut-down |  | . . . . . . . . . . . . . . . . . . . . . | 4.2-23
| 17.5 | Basic Troubleshooting | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-25
18. | Ammonia Analysis by Technicon Autoanalysis |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-26
| 18.1 | Standards and Stock Solutions Preparation . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-27
| 18.2 | Reagent Preparation . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-27
19. | pH Analysis | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-28
| 19.1 | Standard and Reagent Preparation | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-28
| 19.2 | Calibration of pH Meter . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-30
| 19.3 | Pre-Analysis Calibration . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-30
| 19.4 | pH Test 0.01 N HClO4 Solution . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-31
| 19.5 | Analysis of Working Standard . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-31
| 19.6 | Analysis of Filter Extracts | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-32
20. | Atmospheric Species Concentration Calculations . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-32
| 20.1 | Assumptions of the Annular Denuder System . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-32
| 20.2 | Calculation of Air Volume Sampled, Corrected to Standard Conditions . . . . . . . . |  |  | 4.2-33
| 20.3 | Calculations of Concentration Using Results from IC and Technicon
|  | Autoanalysis . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-34
| 20.4 | Calculations Using Results from pH Analysis . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-35
21. | Variations of Annular Denuder System Usage . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-36
22. | Method Safety . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-37
23. | Performance Criteria and Quality Assurance (QA) . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-37
| 23.1 | Standard Operating Procedures (SOPs) . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-38
| 23.2 | QA Program | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-38
24. | References . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-38
iv
Chapter IO-4
Atmospheric Basic & Acidic Constituents
Method IO-4.2
DETERMINATION OF REACTIVE ACIDIC AND BASIC
GASES AND STRONG ACIDITY OF
ATMOSPHERIC FINE PARTICLES (
1. Scope
1.1 | The quantitative measurement of reactive acidic and basic gases and strong acidity of atmospheric fine
particles in ambient air using annual denuder technology is described in this method. |  | The difference between
Inorganic Compendium Methods IO-4.1 and IO-4.2 is that the latter accounts for possible interference from
the dissociation of ammonium nitrate aerosol from particles collected on the filter by two mechanisms:
NH4NO3 = NH3(g) + HNO3(g)
NH4NO3 + NH4HSO 4 = (NH4) 2SO 4 + HNO3
Consequently, an accurate and quantitative value for determining strong acidity of atmospheric fine
particle is calculated.
1.2 The chemical species that can be determined by this method are gaseous SO2 , HNO 2 , HNO 3 , and NH 3 and
particulate SO=4, NO 3- , NH +4, and H +. | Detection and quantitation limits are given in Table 1.
1.3 | This method is a composite of methodologies developed by the U. S. Environmental Protection Agency
(EPA), Harvard University and the CNR Laboratories (Italy). A number of air pollution studies in Italy, United
States, Canada, Mexico, Germany, Austria, and Spain, and in public health services, epidemiology, and
environmental research centers have used this method.
1.4 The equipment described herein is used to measure acidic and basic gases and strong acidity of atmospheric
fine particles contained in ambient air. The methodology originally was developed for monitoring regional-scale
acidic and basic gases and strong acidity of atmospheric fine particle in support of EPA's field programs
involving the Integrated Air Cancer Research Program and the Acid Deposition Network. |  | Similarly, the
methodology has been used to characterize the urban haze in Denver, Houston, and Los Angeles.
1.5 | The techniques, procedures, equipment, and other specifications comprising this method are derived and
composited from those actually used by contributing research organizations and, therefore, are known to be
serviceable and effective. | At this stage, this method is a unified, consensus, tentative, draft method intended
for further application and testing. Users should be advised that the method has not yet been adequately tested,
optimized, or standardized. | Many of the specifications have been initially established by technical judgement
but have not been subjected to ruggedness testing. | In some cases, alternative techniques, equipment, or
specifications may be acceptable or superior. | In applying the method, users are encouraged to consider
alternatives, with the understanding that they should test any such alternatives to determine their adequacy and
confirm and document their advantages. Information and comments are solicited on improvements, alternative
equipment, techniques, specifications, performance, or any other aspect of the method.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-1
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
2. Applicable Documents
2.1 | ASTM Standards
• D1356 Definitions of Terms Related to Atmospheric Sampling and Analysis.
2.2 | Other Documents
• Ambient Air Studies (1-14).
• U. S. EPA Acid Aerosol Document (15).
3. Summary of Method
3.1 | The annular denuder system (ADS) consists of (1) an inlet with an impactor or cyclone preseparator
designed to remove all particles with a diameter (aerodynamic) of 2.5 μm or greater, (2) annular denuders to
quantitate acidic and basic gases, and (3) a filter pack for atmospheric acidity and particles. |  |  | In operation,
ambient air is drawn through an elutriator-accelerator jet assembly, an impactor frit and coupler assembly, and
past glass denuder walls that have been etched and coated with chemicals that absorb the gaseous species of
interest. | The remaining air stream is then filtered through Teflon® and Nylasorb® membrane filters. Teflon®
and nylon membrane filters are used to capture ammonium and nitrate aerosol and sulfate particulate matter.
Nitric acid and sulfur dioxide will also be collected by the nylon filter but these measurements are treated as
interference. The ADS is illustrated in Figure 1. The field sampling box with the ADS and pump-timer system
is shown in Figure 2.
3.2 | Following each run, the ADS assembly is removed from its field housing, its ends are capped, and it is
brought back to the laboratory. In the laboratory, the assembly pieces are uncoupled and capped. The annular
denuders are extracted with 5 mL of deionized water. |  |  | The extracted solutions are subsequently analyzed for
ions corresponding to the collected gaseous species (see Figure 1). The filters are placed into filter bottles where
5 or 10 mL of the ion chromatographic (IC) eluent are pipetted into each filter bottle with the filters faced
downward and completely covered by |  | the eluent. | The filter bottle is capped and put in an ultrasonic bath for
30 min. | The bottles are stored in a clean refrigerator at 5EC until analysis.
3.3 | The analysis of anion and cation concentrations collected by the denuders and filter pack is typically
performed by ion chromatographic and Technicon® colorimeter autoanalytic procedures. The H + concentration
of extracts from the Teflon® filter downstream of the denuders is performed by pH measurements using
commercially available pH meters calibrated with standards.
4. Significance
4.1 | Reactive acidic (SO2 | and HNO 3 ) and basic (NH3 ) gases and particles are found in the atmosphere as a
result of emissions from a variety of fossil fuel combustion sources, including industrial and commercial
facilities, industrial processes, etc. Measurements of these chemical species are currently being used in a range
of environmental studies, such as in (1) epidemiological programs to assess the impact of acid aerosols on
respiratory impairment, (2) receptor modeling to determine the origin of particles that impact EPA's PM2.5 |  |  | or
Page 4.2-2 |  | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
PM10 | air particulate standard, (3) assessment of the impact of particulate nitrate and sulfate on visibility, and
(4) quantification of the impact of acidic and basic air pollutants on issues related to acid rain.
4.2 Unique features of the annular denuder that separate it from other established monitoring methods are the
elimination of sampling artifacts due to interaction between the collected gases and particles and the preservation
of the samples for subsequent analysis. This preservation is accomplished by removing NH3 in the gas stream
by the citric acid coated denuder and reducing the probability of the particulate sulfate (SO4=) captured by the
filter pack from being neutralized to ammonium sulfate [(NH4 ) 2 SO4 ]. |  | If NH 3 | is not extracted from the gas
stream prior to filtration, particulate sulfate and gaseous sulfur dioxide correction would be required for accurate
measurements. |  | These biases in its configuration and analytical methodology are addressed in this method.
5. Definitions
Definitions used in this document and any user prepared Standard Operating Procedures (SOPs) should be
consistent with ASTM D1356. All abbreviations and symbols are defined within this document at the point
of use.
5.1 | Particulate Mass. | A generic classification in which no distinction is made on the basis of origin, physical
state, and range of particle size. (The term "particulate" is an adjective, but it is commonly used incorrectly as
a noun.)
5.2 | Primary Particles (or Primary Aerosols). | Dispersion aerosols formed from particles that are emitted
directly into the air and that do not change form in the atmosphere. |  | Examples include windblown dust and
ocean salt spray.
5.3 | Secondary Particles (or Secondary Aerosols). | Dispersion aerosols that form in the atmosphere as a
result of chemical reactions, often involving gases. A typical example is sulfate ions produced by photochemical
oxidation of SO2 .
5.4 | Particle. | Any object having definite physical boundaries in all directions, without any limit with respect
to size. | In practice, | the particle size range of interest is used to define "particle." | In atmospheric sciences,
"particle" usually means a solid or liquid subdivision of matter that has dimensions greater than molecular radii
(~10 nm); there is also not a firm upper limit, but in practice it rarely exceeds 1 mm.
5.5 | Aerosol. | A disperse system with a gas-phase medium and a solid or liquid disperse phase. | Often,
however, individual workers modify the definition of "aerosol" by arbitrarily requiring limits on individual
particle motion or surface-to-volume ratio. |  | Aerosols are formed by (1) the suspension of particles due to
grinding or atomization or (2) the condensation of supersaturated vapors.
5.6 | Coarse and Fine Particles. | Coarse particles are those with diameters greater than 2.5 μm but less than
10 μm; fine particles are those with diameters less than 2.5 μm. |  | These two fractions are usually defined in
terms of the separation diameter of a sampler.
[Note: Separation diameters other than 2.5 μm have been used.]
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-3
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
5.7 | Annular. | Refers to rotating to, or forming a ring. | In the annular denuder sampler, the annular refers to
the cylinder to which coating is applied to the interior parallel planes to remove gaseous pollutants by diffusion
chemistry.
5.8 | Denuder. | To strip away the process gaseous pollutants from the gas stream.
5.9 | Equivalent Weight. | The equivalent weight, or combining weight, of a compound or ion is its formula
weight divided by the number of replaceable hydrogen atoms.
5.10 | Normal Solution. | Solution that contains a gram-equivalent weight of solute in a liter of solution.
6. Factors Affecting Denuder Efficiency
6.1 | Operation below 20% relative humidity (RH) may result in less than quantitative collection of SO2 .
Atmospheric water vapor in concentrations above 30% RH has been shown not to be an interferant for SO2
collection.
6.2 | Studies are being conducted to identify interferents, and calculations are being developed to correct the
measurements obtained by the annular denuder system for identifiable interferents. For example, the presence
of ozone (O 3 ) is known to oxidize nitrous acid (HNO2 ) to nitric acid (HNO3 ); therefore, HNO 2 |  |  | measurements
are often underestimates. Calculations have been developed to adjust for this oxidation process and to provide
more accurate estimations of HNO2 |  | concentrations in the atmosphere.
6.3 | Other studies include the possible chemical reactions (organic and inorganic) that may occur with selected
coating solutions that interfere with the accurate measurement of the chemical species of interest.
6.4 | The efficiency of impactor collection decreases when the impactor surface is loaded. |  | The average
operational time before such loading occurs has not been determined.
7. Apparatus
[Note: | This method was developed using the annular denuder system produced by University Research
Glassware, 116 S. Merritt Mill Road, Chapel Hill, NC 27516, (919) 942-2753, as a guideline. |  |  | EPA has
experience in use of this equipment during various field monitoring programs over the last several years.
Other manufacturers' equipment should work as well. Denuder systems are also available from Rupprecht
and Patashnick Co., Inc., 25 Corporate Circle, Albany, NY 12203, (518) 452-0065 and Ogawa & Company,
USA, Inc., 1230 S.E. 7th Avenue, Pompano Beach, Florida 33060, (305) 781-6223. However, modifications
to these procedures may be necessary if another commercially available sampler is selected.]
7.1 | Sampling
| 7.1.1 Elutriator and Acceleration Jet Assembly (see Figure 3). |  | Under normal sampling conditions, the
elutriator or entry tube is made of either Teflon®-coated glass or aluminum. When using glass, the accelerator
jet assembly is fixed onto the elutriator, and the internal surfaces of the entire assembly are coated with
Teflon®. | When aluminum is used, the accelerator jet assembly is removable. |  | The jet is made of Teflon® or
Page 4.2-4 |  | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
polyethylene, and the jet support is made of aluminum. |  | Again, all internal surfaces are coated with Teflon®.
Both assemblies are available with 2, 3 and 4 mm inside diameter jets (nozzles).
7.1.2 Teflon® Impactor Support Pin and Impactor Frit Support Tools (see Figure 4). Made of either
Teflon® or polyethylene and used to aid in assembling, removing, coating and cleaning the impactor frit.
7.1.3 | Impactor Frit and Coupler Assembly (see Figure 5). | The impactor frit is 10 mm x 3 mm and is
available with a porosity range of 10-20 μm. |  | The frits should be made of porous ceramic material or fritted
stainless steel. | Before use, the impactor frit surface is coated with a Dow Corning 660 oil and toluene solution
and sits in a Teflon® seat support fixed within the coupler. |  | The coupler is made of thermoplastic and has
Teflon® clad sealing "0"-rings that are located on both sides of the seat support inside the coupler. |  | The
couplers are composed of two free moving female threads that house the support tools when assembling and
removing the impactor frit, and couple the denuders when sampling. Arrows on the metal band hold the female
threads together. | These arrows should be pointing in the direction of air flow (see Figure 1) when the ADS is
assembled.
[Note: In situations when there are substantial high concentrations of coarse particles (>2.5 μm), a Teflon®-
coated aluminum cyclone should be used in place of the acceleration jet and impactor assembly, as illustrated
in Figure 6.]
The cyclone is made of Teflon®-coated stainless steel. |  | The location of the cyclone with respect to the
denuder, heated enclosure, and meter box is illustrated in Figure 6.
7.1.4 | Annular Denuder (see Figure 7). | The original denuders (single channel denuders) consist of two
concentric glass tubes (see Figure 4). | The tubes create a 1 mm orifice, which allows the air sample to pass
through. | The inner tube is inset 25 mm from one end of the outer tube; this end is called the flow straightener
end. | The other end of the inner tube is flush with the end of the outer tube. | Both ends of the inner tube are
sealed. | In this configuration, the glass surfaces facing the orifice are etched to provide greater surface area for
the coating. | Three types of denuders are available. | One is the older version that accommodates the impactor
support pin assembly and can only be the first denuder in sequence. |  | It is available in glass with the impactor
support holder made of glass and the impactor support pin assembly made of Teflon®. The denuder is 265 mm
long with size #30 threads for coupling. | It is available with flow straighteners at both ends; however, most
denuders in use today only have one flow straightener end. |  | The second most recent denuder version, which
can be used as any denuder in sequence, is available in glass with only one flow straightener end. |  | It is 242 mm
long and has size #30 threads. | Finally, the third denuder (multi-channel denuder) design involves two inner
concentric glass tubes (1 mm separation) positioned around a solid center glass rod, as illustrated in Figure 7.
Once again, the glass surfaces are etched to provide greater surface area for the coating. |  | The inner glass tubes
and coater rod are inset 25 mm from one end of the outer Teflon®-coated stainless steel tube to serve as the
flow straightener end. Since 1992, the multi-channel denuder has been evaluated and used extensively by EPA
in place of the single channel, all glass denuder. |  | The stainless steel sheath multi-channel denuder, while more
expensive, is easier to handle during coating, extraction procedures, and in field applications. |  | Additional
channels in the denuder increases capacity of the denuder, thus minimizing breakthrough.
This procedure uses two denuders in series. The first denuder, Na 2CO3 coated, is used to trap out HNO 3
and SO 2. The Na2CO3 denuder can also trap out HCl, HNO 2, HF, and some organic acids. However, special
care needs to be taken in the preparation of the denuder coating material and extraction procedures to
quantitate HCl. | This method does not quantitate HCl, HNO2, HF, or other organic acids retained on the
Na2CO3 denuder.
The second denuder is coated with citric acid for trapping NH 3. |  | This procedure does cover the
quantitation of NH3.
All denuder types should be equipped with thermoplastic or polyethylene caps when purchased.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-5
Method IO-4.2 | [10 empty] | Chapter IO-4
Acidic/Basic Constituents | [8 empty] | Atmospheric Acidic
7.1.5 | Caps for Annular Denuder (see Figure 2). |  |  |  |  | Caps are made of either polyethylene or thermoplastic
and are used in the coating and drying processes and for storage and shipment. The thermo-plastic caps include
a removable Teflon® seal plate when purchased. |  |  |  |  | Repeated reuse of these types of caps have caused some
contamination due to improper cleaning of the cap and Teflon® seal plate (i.e., fluid tends to be trapped under
the seal plate). | The polyethylene caps are not equipped with seal plates. | [6 empty] | Polyethylene caps tend to dry faster
and seal better than the thermoplastic caps.
7.1.6 | Annular Denuder Couplers (see Figure 4). |  |  |  |  | The couplers should be made of thermoplastic and
equipped with Teflon® "O"-rings that sandwich a silicone rubber ring on three sides. | [8 empty] | This design provides
elasticity for better sealing under extremely cold temperature conditions in which Teflon® does not give. | [10 empty] | Two
types of couplers are available. |  | In the older version, the couplers have removable seal rings. | [8 empty] | Problems with
denuder breakage and leakage due to improper threading of the couplers with the denuders led to the
development of a second type of coupler. |  |  |  | The new couplers are equipped with permanent seal rings that
provide more even threading and a better seal when coupled. | [6 empty] | Some couplers have built-in flow-straighteners.
The couplers are used to couple the annular denuders and for coupling the last denuder with the filter pack.
7.1.7 Filter Pack Assembly (see Figure 8). The filters are supported by stainless steel porous screens and
housed in a polyethylene filter ring housing. The Teflon® filter ring housing directly follows the Teflon® filter
housing inlet component. |  | The "nylon" filter ring housing follows the Teflon® filter ring housing and sits on a
Teflon® "O"-ring, which seals the filter ring housing components to the filter housing outlet component. | [10 empty] | The
filter housing outlet component is aluminum and accommodates a polyethylene screw sleeve that seals the filter
pack assembly. | The sleeve is available in different lengths to accommodate up to four filter ring housing units.
A stainless | steel | "Quick-Release" | plug | screws | into | the | aluminum | outlet | component | for | connecting | the
pump-timer to the filter pack assembly. It is equipped with an orange "dust cover" (male plug) upon purchase.
7.1.8 | Drying Manifold Assembly (see Figure 9). |  |  |  |  | Made of pyrex and is available to accommodate as
many as four drying denuders. |  | The denuders are attached to the manifold with back-to-back Bakalite bored
caps. | The bored caps are connected with a Teflon® connector ring. Air is pushed through an air dryer/ cleaner
bottle made of 2 1⁄2" heavy wall pyrex that contains silica gel, calcium sulfate, and activated charcoal (not
available with assembly). |  | The tubing that connects the dryer/cleaner bottle to the drying manifold should be
secured at each cap with either Teflon® washers or Teflon® washers coupled with Teflon® hose barbs.
A mixture of activated carbon/drying agent/sodium carbonate flux (~50 g) is contained in the
dryer/cleaner bottle to scrub contaminants from the supply air source.
7.1.9 | Vacuum Tubing. | Low density polyethylene tubing, 3/8" diameter for distances of less than 50 ft.,
1⁄2" diameter for distances greater than 50 ft. |  |  |  | Since this tubing is used downstream from the sampler, similar
sized tubing or pipe of any material may be substituted. | [5 empty] | The tubing must have sufficient strength to avoid
collapsing under vacuum [Fisher-Scientific, 711 Forbes Ave., Pittsburgh, PA 15219, (412-787-6322)].
7.1.10 | Tube Fitting. | Compression fittings (Swagelok®, Gyrolok® or equivalent) to connect vacuum
tubing (above) to an NPT female connector or filter holder and connect vacuum tubing to fitting on differential
flow controller. | The fittings may be constructed of any material since they are downstream of the sampler
[Fisher-Scientific, 711 Forbes Ave., Pittsburgh, PA 15219, (412-787-6322)].
7.1.11 | Annular Denuder System Sampling Case (see Figure 10). | [6 empty] | Made of a "high-impact" plastic and
insulated with polyurethane. |  | It is 4 ft long by 6" wide and 6" deep. |  |  |  |  | Two heater units, a fan blower, and an air
outlet are located in the lid of the housing. |  |  |  | Also, located on the lid are the automatic and manual control
switches and a 12-V power supply outlet for the heater and fan. | [6 empty] | The bottom of the box houses the ADS. |  |  |  | The
elutriator end of the ADS protrudes through one end of the box, while the denuders are supported in the box
by chrome plated spring clips. |  | If the Teflon®-coated aluminum cyclone is used to remove coarse particles, it
is also housed in the heated sampling box, with the elutriator end protruding through the sampling box. | [11 empty] | A
vacuum plug known as a "quick-release" coupler is linked to the filter pack of the ADS. This plug connects the
ADS to 1 1/4" | Teflon® rubber "clad" shrink tubing that exhausts the air stream to the ambient air. | [9 empty] | The box
is sledge hammer proof.
Page 4.2-6 |  | Compendium of Methods for Inorganic Air Pollutants | [8 empty] | June 1999
Chapter IO-4 |  |  |  |  | Method IO-4.2
Atmospheric Acidic |  |  |  | Acidic/Basic Constituents
7.1.12 Annular Denuder Field-to-Lab Case (Optional). Annular Denuder Field-to-Lab Case (Optional).
The field-to-lab case is made of rigid plastic and insulated with polyurethane. It is made to be hand carried, not
shipped, and is used to transport four total annular denuder systems, each consisting of either three annular
denuder sections or two annular denuder sections and one denuder-impactor assembly. The systems are packed
already assembled and capped; they are ready for sampling or sample analysis. The case has a carrying handle,
a lock, and three latches and is equipped with two keys.
| 7.1.13 | Annular Denuder Shipping Case (Optional). |  | Made of formica, backed with plywood and
insulated with polyurethane. The corners are reinforced with metal. It is made to withstand shipping by truck,
UPS, and Federal Express. Each case is stackable and lockable and has a carrying handle. Seven total annular
denuder systems can be packed in the case, provided each system contains four denuders each. |  |  |  |  | The systems
can consist of either three denuders (242-mm long) and one denuder-impactor assembly (265 mm long) or four
denuders (242-mm long). |  |  | Each component of the system is packed in its own storage compartment. |  | The
personal sampler assemblies can also be placed and shipped in this case.
| 7.1.14 | Differential Flow Controller (Pump). |  | Pumps air through the sampler at a fixed rate of between
5 and 20 standard L/min (typically 10 L/min) with a precision of ±5% over the range of 25 to 250 mm Hg
vacuum.
| 7.1.15 | Dry Gas Meter (DGM). |  | Pulls 10 L of gas per revolution [Fisher-Scientific, 711 Forbes Ave.,
Pittsburgh, PA 15219, (412-787-6322)].
7.2 | Analysis
| 7.2.1 | Ion Chromatograph. | A chromatograph equipped with the appropriate anion and cation exchange
resin filled separator and suppressor columns and conductivity detector for measuring acidic (SO2 , HNO 2 and
HNO 3 ) and basic (NH3 ) ions in solution (i.e. denuder and filter extracts) [Dionex Corp., 1228 Titan Way,
Sunnyvale, CA 94086, (408-737-0700)].
| 7.2.2 Technicon Colorimeter Autoanalyzer. |  |  | Colorimetic analyzer able to detect specific ions of interest
in aqueous extracts [Technicon Industrial Systems Corp., 511 Benedict Ave., Tarrytown, NY, 10591-5097,
(800-431-1970)].
| 7.2.3 | pH Meter. | A pH or pH/ion meter with an "integral" automatic temperature compensation and
calibrated with standard buffers (pH 4 and 7), including 2 and 4 mL analysis cups (Orion and other vendors).
| 7.2.4 | Polyethylene Bottles with Polyethylene Screw Caps. |  | 50 mL and 100 mL. | Used for storage of
coating solutions, best source.
| 7.2.5 Erlenmeyer Flasks. |  | 250 mL and 2 L borosilicate glass or polyethylene flasks with calibration, best
source.
| 7.2.6 Graduated Cylinders. |  | 10 mL and 100 mL borosilicate glass or polyethylene cylinders, best source.
| 7.2.7 | Pipets. | Class A 5 mL and 10 mL borosilicate glass pipettes or automatic pipettes. |  | Calibrated "to
deliver," best source.
| 7.2.8 | Pipet Bulb. | Made of natural rubber. | Recommended to meet OSHA requirements, best source.
| 7.2.9 Micropipettes. |  | Recommended 50 μL. Calibrated "to contain," borosilicate glass micropipette, best
source.
7.2.10 Forceps. Recommended dressing forceps made of stainless steel or chrome-plated steel and without
serrations. |  | Used for handling filters (Millipore).
| 7.2.11 | Stopwatch. | Used for measuring flow rate of gas stream through DGM, best source.
| 7.2.12 | Ultrasonic | Cleaner. | Used for filter extractions and parts cleaning. | Most are temperature
controlled. |  | Temperature should be controlled during extraction at 65EC [Cole-Palmer Instrument Co., 7425
N. Oak Park Ave., Chicago, IL 60648, (980-323-4340)].
| 7.2.13 | Clean Air Hood. | Closed air hood with ammonia free air circulation. | Used for Teflon® filter
extraction for pH analysis, best source.
June 1999 |  |  | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-7
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
8. Reagents and Materials
8.1 | Teflon® Filters. | Zefluor® (PTFE) membrane filters 47-mm diameter with a 2 μm pore size. |  | Only one
side is Teflon®-coated; this side should face the air stream [Gelman Sciences, 600 S. Wagner Rd., Ann Arbor,
MI, 48106, (800-521-1520)].
8.2 | Nylasorb® Filters. | Membrane filters 47-mm diameter with a 1 μm pore size. | These filters are specially
prepared and batch analyzed for low NO3 G background levels. |  |  | If other brands of nylon membrane filters are
used, they should be batch analyzed to ensure low and replicable levels of |  |  | NO3 G [Gelman Sciences, 600 S.
Wagner Rd., Ann Arbor, MI, 48106 (800-521-1520)].
8.3 | Denuder Extract Storage Vials. | 30 mL (1 oz) screw-cap polyethylene sampling vials (Nalgene or
equivalent). |  | Allow eight per sample for each sampling period, best source.
8.4 Filter Extract Storage Vials. |  | 25 mL polyethylene vials (Nalgene or equivalent). Allow two vials for each
sampling period, best source.
8.5 | IC Analysis Vials and Caps. | Available in 5 mL and 0.5 mL and are made of polypropylene. |  | The filter
caps are made of plastic and contain a Teflon® filter through which the sample is extracted for analysis. Both
the vials and filter caps should be disposable, best source.
8.6 | Labels. | Adhesive, for sample vials, best source.
8.7 | Parafilm. | Used for covering flasks and pH cups during pH analysis, best source.
8.8 | Kimwipes® and Kay-dry Towels. | Used for cleaning sampling apparatus and analysis equipment, best
source.
8.9 | Stoppers. | Cork or polyethylene, best source.
8.10 | Sodium Carbonate (Na 2 CO 3 ). | ACS reagent grade, best source.
8.11 | Sodium Chloride (NaCl). | ACS reagent grade, best source.
8.12 | Toluene. | ACS reagent grade, best source.
8.13 | Glycerol (Glycerine - CH2 OHCHOHCH2 OH). |  | ACS reagent grade, best source.
8.14 | Citric Acid [Monohydrate - HOC (CH 2 CO) OH] 2 COOH : H2 O. |  | ACS reagent grade, best source.
8.15 | Hydrogen Peroxide (H 2 O2 ). | ACS reagent grade, best source.
8.16 | Ethanol (C2 H5 OH). | ACS reagent grade, best source.
8.17 | Sulfuric Acid (H2 SO4 ). | ACS reagent grade, best source.
Page 4.2-8 |  | Compendium of Methods for Inorganic Air Pollutants |  | June 1999
Chapter IO-4 | [11 empty] | Method IO-4.2
Atmospheric Acidic | [9 empty] | Acidic/Basic Constituents
8.18 | Potassium Chloride (KCl). |  |  |  | ACS reagent grade, best source.
8.19 | Perchloric Acid (HClO4 ). |  |  | ACS reagent grade (60-62%), best source.
8.20 | Distilled Deionized Water (DDW). |  |  |  |  | ASTM Type I water.
8.21 | pH Buffers. | Standard buffers 4.00 and 7.00 for internal calibration of pH meter, best source.
8.22 | Silica Gel. | ACS reagent grade (indicating type), best source.
8.23 | Sodium Bromide (NaBr). |  |  | ACS reagent grade, best source.
8.24 | Activated Charcoal. |  | ACS reagent grade, best source.
8.25 Balance. Electronic analytical with internal calibration weights and enclosed weighing chamber. Precision
of 0.1 mg [Fisher-Scientific, 711 Forbes Ave., Pittsburgh, PA, 15219, (412-787-6322)].
8.26 Gloves. | Polyethylene disposable. Used for impactor frit assembly and filter pack assembly, best source.
8.27 | Dow Corning High Temperature Vacuum Oil. | [6 empty] | Dow Corning 660 oil used for impactor frit coating
solution, best source.
8.28 | Zero Air. | A supply of compressed clean air, free from particles, oil, NO, NO2 , SO2 , HNO 3 , and HNO 2 .
The supply may be either from a commercial cylinder or generated on site, best source.
8.29 | Ion Chromatographic (IC) Eluent Solution. | [6 empty] | For extracting filters. |  | This solution should be the same
eluent as used for the ion chromatographic analysis of the filters. | [8 empty] | If the filter analysis is not to be performed
by ion chromatography, a slightly basic solution (e.g., 0.003 N NaOH or sodium carbonate/bicarbonate) should
be used to extract the Nylasorb® filter, while the Teflon® filter should be extracted with double-distilled
deionized water (DDW).
9. Preparation of Coating and Extraction Reagents
9.1 Impactor Frit Coating Solution Preparation. | [7 empty] | Weigh 1 g of silicone oil (Dow Corning high temperature
660 oil) and place in a 100 mL polyethylene storage bottle. | [7 empty] | Add 100 mL of toluene. |  |  | Mix thoroughly, close
container, and store at room temperature. | [5 empty] | (WARNING - FLAMMABLE LIQUID).
9.2 | Impactor Frit Extraction Solution Preparation. | [6 empty] | Add 100 mL of IC eluent to a clean polyethylene
storage | container. | Pipette | 5 | mL | of | ethanol | into | container. | Mix | thoroughly. | Store, | covered | at | room
temperature.
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | [9 empty] | Page 4.2-9
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
9.3 | Annular Denuder Coating Solutions Preparation
[Note: Different coatings may be used depending on the chemical species of interest.]
9.3.1 Na 2 CO3 Coating Solution. Clean a 100 mL polyethylene storage vial and let dry at room temperature.
Measure 50 mL of ethanol (WARNING - TOXIC, FLAMMABLE LIQUID) with a graduated cylinder and pour
into vial. | Measure 50 mL of DDW with a graduated cylinder and add to vial. |  |  | Weigh 1 g of glycerol and add
to DDW. Weigh 1 g of Na2 CO3 | and add to vial. |  | Mix thoroughly, solution may fizz; wait for fizzing to stop
before sealing vial. | Store at room temperature.
9.3.2 | Citric Acid Coating Solution. | Clean a 100 mL polyethylene storage vial and let dry at room
temperature. | Measure 50 mL of ethanol (WARNING - TOXIC, FLAMMABLE LIQUID) with a graduated
cylinder and pour into vial. | Weigh 0.5 g | of | citric acid and add to a vial. | Add enough glycerol to vial to make
it a 1% solution (0.5 mL). | Mix thoroughly; store, covered at room temperature.
10. Elutriator and Acceleration Jet (Inlet) Assembly
[Note: The all-glass configuration is shown in Figure 6A.]
10.1 The internal walls of the elutriator and jet assembly are coated with Teflon® to prevent losses of reactive
species (SO 2 , HNO 3 , NH 3 ) during sampling. The elutriator prevents water and large particles from entering the
inlet and thus extends the life of the impaction surface located immediately downstream of this assembly.
10.2 | An aluminum version of this inlet is shown in Figure 3b. |  | All inner surfaces of the aluminum unit are
Teflon®-coated. | The main difference between the all glass and the aluminum inlet is the jet component of the
aluminum inlet is replaceable, as shown in Figure 3b. |  |  | The jet component is made of either Teflon® or
polyethylene and is available in various diameters as needed to accommodate selected sample flow rates. |  |  |  | The
jet may be replaced using the tool shown in Figure 3b. |  |  | The jet diameter for a sample flow rate of 10 L/min is
3.33 mm. | At this flow rate, the inlet has an aerodynamic cutpoint of 2.5 μm. |  |  | If a different flow rate is used,
the jet diameter must be changed to retain an aerodynamic cutpoint to 2.5 μm. |  |  |  | Table 2 contains the jet
diameters and Reynolds number to maintain an aerodynamic of 2.5 μm cutpoint at different flow rates between
1 and 20 L/min.
[Note: | If the sampling area has substantial concentrations of coarse particles (>2.5 μm), the user may
replace the acceleration jet and impactor assembly with the Teflon®-coated aluminum cyclone.]
11. Impactor Frit Preparation and Installation
11.1 | Impactor Frit Installation
11.1.1 Impactor-Coupler. The impactor-coupler assembly (shown in Figure 12) is composed of two parts:
the replaceable impactor frit and the coupler-impactor housing seat. |  |  | The impactor surface is a porous ceramic
or porous stainless steel frit, 10 mm x 3 mm. |  |  | This frit is inserted into the coupler-impactor housing using the
tools shown in Figure 12. | The in-tool must be completely screwed in behind the impactor seat before the frit
is pressed into place. | Press the impactor frit gently but firmly into the seat of the impactor housing with your
Page 4.2-10 | Compendium of Methods for Inorganic Air Pollutants |  |  | June 1999
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
clean gloved finger. | The impactor should fit into the housing so that it does not protrude above the seat. |  | The
impactor frit has a slight bevel. |  | The narrow surface should be inserted into the impactor seat.
11.1.2 | Impactor-Denuder. | The impactor-denuder assembly (shown in Figure 4) is of three parts: | the
replaceable impactor frit, the impactor seat support pin, and the annular denuder impactor-pin support. |  |  | The
impactor frit is the same as described in Section 11.1.1 and is inserted into the impactor seat support pin. |  |  | The
impactor support pin can be hand-held while inserting the frit or it can be placed upright into the stainless steel
frit holder #3 (see Figure 11). | Press the support pin into the denuder pin support. |  | The pin is grooved and has
a viton "O"-ring to keep the pin snug in the denuder support during cold weather use (Teflon® tends to shrink
at low temperatures). | The support pin is removed by using the removal tool shown in Figure 4.
11.2 | Impactor Frit Preparation
With the impactor frit in the impactor seat of either the coupler (see Figure 12) or the Teflon® impactor seat
support pin that fits into the first denuder (see Figure 4), pipette 50 μL of the toluene-660 oil coating solution
onto the impactor frit surface and allow to dry at room temperature. Cap both sides of the coupler impactor
or denuder-impactor until use.
12. Filter Pack Preparation and Assembly
[Note: | Any number of filters can be used depending on the target species of interest. |  | The configuration
referred to in this section does not collect NH +4 |  | .]
12.1 With clean gloves, disassemble the filter pack (see Figure 8) by unscrewing the large outer Teflon® collar
(sleeve) from the aluminum filter housing outlet component.
[Note: Remove the red Bakelite® cap first. Lay the pieces out on clean Kimwipes®.]
12.2 | Lay a clean Teflon® filter ring housing, with its large opening face up, on a clean Kimwipe®. |  | Place a
clean stainless steel screen in the filter ring housing.
12.3 Using clean filter forceps, place a Nylasorb® nylon filter on the screen. Insert a second filter ring housing
on top of the nylon filter with its large opening face up. |  |  | This design forms a "sandwich" with the nylon filter
held between the two filter ring housings.
12.4 | Place another clean screen on the second filter ring housing. |  | Using clean filter forceps, place a Teflon®
filter on the screen.
[Note: If a Teflasorb® Teflon® filter is used, be sure to place the Teflon®-coated side, not the webbed side,
toward the air stream. If the webbed side is facing the air stream, SO 4= extraction from the filters may be
inefficient.]
12.5 Place the Teflon® filter housing inlet component (see Figure 11) on top of the Teflon® filter, which forms
another "sandwich" with the Teflon® filter held between the second filter ring housing and the housing inlet
component. The housing inlet component connects the filter pack assembly to the last annular denuder through
a thermoplastic coupler. | Be careful not to twist the filterpack components, or damage will occur to the filters.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-11
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
12.6 | Lay the aluminum filter housing outlet component, with its large opening face up, on a clean Kimwipe®.
12.7 | Insert the filter ring sandwiches (prepared in Sections 12.1 through 12.5) with the filter housing inlet
component extending upward, in the aluminum filter base. |  | Place the large outer Teflon® sleeve over the filter
sandwich and screw onto the aluminum filter base. |  | DO NOT OVERTIGHTEN! | DO NOT TWIST FILTER
PACK COMPONENTS!
12.8 | Install the "Quick-Release" plug into the filter outlet component. | DO NOT OVERTIGHTEN!
12.9 | Install the red Bakalite® cap onto the filter inlet component and the orange dust cover onto the
Quick-Release plug until ready to attach denuders.
13. Annular Denuder System Preparation
All new annular denuder parts obtained from suppliers should be cleaned with a dilute soap solution. The
parts should then be thoroughly rinsed in DDW and allowed to dry at room temperature.
13.1 | Annular Denuder Coating Procedure
[Note: If the first denuder holds the impactor, a blank Teflon® impactor support pin should be installed in
the pin support holder before the coating procedure.]
13.1.1 | Cap the end of the denuder that has the inner tube flush to the outer tube and set denuder upright
on the capped end. | For the denuders with flow-straighteners at both ends, either end can be capped. |  | Measure
10 mL of the appropriate coating solution into a graduated cylinder. Pipette the 10 mL into the flow-straightener
end of the upright capped annular denuder.
13.1.2 | Cap the open end of the denuder, and holding horizontally, rotate the denuder to distribute the
coating solution evenly.
13.1.3 | Remove cap from flow-straightener end of denuder and decant excess coating solution into a clean
denuder extract storage bottle labeled "denuder blank." |  | Bottle label should include denuder number, coating
solution and date.
13.1.4 | Repeat this procedure with each denuder; label the denuders and bottles appropriately.
13.2 | Annular Denuder Drying Procedure (see Figure 9)
[Note: As denuders dry, they change from translucent to a frosted appearance. Denuders are dry when they
become uniformly frosted.]
13.2.1 Adjust drying train and manifold clean air flow to 2 to 3 L/min. Close toggle valve controlling clean
air flow through manifold before attaching denuders.
13.2.2 Attach flow-straightener end to drying manifold port at the back-to-back bored caps (see Figure 9).
13.2.3 | Open toggle valve and allow clean air to flow through the tube for several minutes.
13.2.4 | Close toggle valve and reverse ends of tubes attached to manifold.
13.2.5 | When an even frosted appearance is achieved, remove tubes from manifold, cap both ends with
clean caps, and store until ready for use. | Turn off air to drying manifold.
Page 4.2-12 | Compendium of Methods for Inorganic Air Pollutants |  | June 1999
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
13.3 | Annular Denuder System (ADS) Assembly
[Note: Described herein is an annular denuder system consisting of four denuders in series. Any number of
denuders can be used at the operators' discretion. The denuders should be assembled in such a way that the
flow-straightener end always follows the flush end of the previous denuder, except, if denuders with flow-
straighteners at both ends are used. |  | This type of assembly allows laminar flow conditions to be restored
quickly.]
13.3.1 | Lay the ADS pieces on a clean surface (i.e., Kimwipes®).
13.3.2 | Remove the end caps from the first denuder. | Denuder 1 is coated with Na2 CO3 | and holds the
impactor frit pin support. | Remove the blank impactor support pin. | Gently insert the impactor support pin and
coated frit assembly into the denuder-pin support. |  | If the first denuder does not hold the impactor pin-support,
attach the impactor frit seat equipped coupler assembly to the flow-straightener end of the first denuder.
[Note: DO NOT TIGHTEN! Do not tighten during the following procedure until Section 13.4.8 is reached.]
13.3.3 | Attach a thermoplastic coupler to the opposite denuder end. | Place a Teflon® clad "O"-ring inside
the coupler, if needed.
13.3.4 | Remove the end caps from the second denuder (citric acid coated). | Attach the end with the
flow-straightener section to the denuder-coupler assembly.
13.3.5 | Attach a thermoplastic coupler to the opposite denuder end. | Place a Teflon® clad "O"-ring inside
the coupler, if needed.
13.3.6 | Attach the filter pack inlet to the second denuder coupler assembly.
13.3.7 | Attach the elutriator-acceleration jet assembly to the first denuder-coupler assembly. Tighten very
gently--DO NOT OVERTIGHTEN!
13.3.8 | Tighten the remaining couplers very gently - DO NOT OVERTIGHTEN!
13.3.9 | Cap elutriator with orange dust cover until use.
[Note: | When collecting and measuring gaseous HNO 3, SO 2, and particulate NO 3- and SO 4= , NH 3 must be
taken out of the gas stream prior to the air stream entering the filter pack. |  | Otherwise, reaction of the
unneutralized sulfate will result. If ammonia (NH 3) and/or H + measurements are not to be analyzed for, then
the use of a citric acid coated denuder is not important. |  | However, with the removal of NH 3, some nitrate
collected on the Teflon® filter will evaporate and be found on the nylon filter.]
13.4 | Laboratory Leak-Check of ADS
[Note: CAUTION--Do not subject the system to sudden pressure changes or filters may tear.]
13.4.1 | Remove the orange dust cap from the impactor opening. | Attach the "Quick-Release" to a pump
module. | Turn on the pump. | Be certain that flow through the ADS occurs by checking the rotameter.
13.4.2 | Briefly cap the elutriator with the orange dust cap. | The flow as indicated on the rotameter should
drop to zero if no leaks exist.
13.4.3 | Disconnect the pump from the ADS at the "Quick-Release" plug. | Cap the "Quick-Release" plug
with an orange dust cover. Turn off the pump. REMEMBER--never overtighten joints or breakage will result.
If the joints can not be sealed with gentle tightening, the Teflon® "O"-rings are worn or defective and must be
replaced.
13.4.4 | Place the assembled sampler in its field-to-lab carrying case for transport to the field.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-13
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
[Note: The ADS joints should be loosened slightly when extreme temperature changes are incurred during
transportation. |  | This precaution will prevent unnecessary breakage or distortion of the ADS components.
Remember to allow the system to adjust to the indoor air temperature before tightening the joints and checking
for leaks.]
14. Sampling
14.1 | Placement of Denuder System
14.1.1 | The placement of the annular denuder system must conform to a consistent set of criteria and
guidance to ensure data comparability and compatibility. |  | A detailed set of monitor siting criteria for ambient
air monitoring and meteorological programs is given in the EPA document Ambient Monitoring Guidelines for
Prevention of Significant Deterioration (PSD), EPA-450/4-87-007, U. S. Environmental Protection Agency
Office of Air Quality Planning and Standards, Research Triangle Park, NC 27711, May 1987.
The site must be away from localized sources of ammonia, such as composting and livestocking
operations, landfills, sewage treatment plants, fertilizer plants and storage facilities, and recently plowed
fertilized fields because aerosol acidity is subject to rapid neutralization by ambient bases.
14.1.2 | A summary of key factors that should be considered as part of the placement of an air quality
monitoring station containing an ADS are:
• Vertical placement above ground;
• Horizontal spacing from obstructions and obstacles;
• Unrestricted air flow; and
• Adequate spacing from roads.
The ADS sampler is mounted on a supported mast pole or tripod. The ADS inlet should be located 2-3
m above ground level. Placing the inlet closer to ground level should be considered only if the surface is flat
and man-made (i.e., not unpaved dirt).
14.1.3 A summary of key criteria associated with these siting factors for air monitoring stations is included
in Table 3.0. | The information included in the table should be used to the extent possible as part of the
monitoring network design to ensure that the monitoring program provides representative and unbiased data.
However, site-specific constraints could make it very difficult to meet all criteria. For example, wooded areas
around a site would make the siting very difficult. |  | The use of the information in Table 3.0, coupled with a
balanced evaluation by an experienced air quality and meteorology specialist, is highly recommended.
14.1.4 In general, for a site with no major obstruction and obstacles, the air sampler intake should be about
2-3 m aboveground. For a site with nearby roadways, however, intake placement should take into account the
effects of road dust re-entrainment and vehicular emissions. |  | In fact, a linear relationship should be established
between the horizontal distance of the sampler intake from the roadway and the aboveground elevation of that
intake. | For any roadway accommodating more than 3,000 vehicles per day, the intake should be between 5
and 25 m from the edge of the nearest traffic lane. |  | It should also be 15 m aboveground for a distance of 5 m
from the nearest traffic lane and 2 m aboveground for a distance of 25 m from the nearest lane. For a roadway
supporting less than 3,000 vehicles per day, the intake should be placed at a distance greater than 5 m from the
edge of the nearest traffic lane and at a height of 2-15 m aboveground.
14.2 | Start-Up
14.2.1 Remove the ADS from its field-to-lab carrying case and load into the field sampling box. |  |  | The ADS
field sampling box is insulated with polyurethane which is configured to hold the ADS without allowing
movement. | Chrome plated spring clips hold the denuders in place. | Automatic and manual control switches
Page 4.2-14 |  | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 | [12 empty] | Method IO-4.2
Atmospheric Acidic | [11 empty] | Acidic/Basic Constituents
allow the sampling box to control the temperature of the ADS. | [7 empty] | The automatic switch should be used when the
ADS is not in use and when the ADS is sampling for extended periods of time without constant supervision to
prevent low temperature or sudden pressure change exposure of the ADS (these types of exposure can cause
leaks to occur, condensation, or the filters to tear). |  |  |  |  | When sampling, the ADS should be kept 1EC above the
ambient temperature to prevent condensation. The sampling box has two connections with the pump timer: the
plastic | suction | hose | connected | with | "Quick-Release" |  | couplers |  |  | and | the | 12-V | power | cord | with | a
"Quick-Disconnect" coupler. |  |  | The power cord remains connected, and the suction hose is disconnected from
the box each time the unit is opened. |  |  |  | Inside the box, the hose is connected to the top of the filter pack with a
"Quick-Release" coupler. |  |  | During sampling the sample box is kept securely closed (see Figure 2).
14.2.2 | Allow the pump to warm up for 20-30 min prior to testing so the pump will provide steady flow
during testing.
14.2.3 To check the Heat/Cool cycles, flip one switch from "AUTO" to "MANUAL" and the other between
"COOL" and "HEAT." |  |  | Check to insure that the fan and heater work, respectively.
14.2.4 With the elutriator still capped, turn on the pump with the switch on the timer. The rotameter should
indicate zero flow. |  | If there is a flow, the assembly pieces need to be recoupled. | [9 empty] | Run leak check for 5-10 s.
Turn off pump and remove elutriator cap. |  |  |  |  | Record leak rate on Field Test Data Sheet (see Figure 13).
14.2.5 | Attach DGM output to elutriator inlet. |  |  |  | Turn on pump. |  |  |  | Record start time on Field Test Data Sheet
(see Figure 13). |  | Using a stopwatch, record the time for 20.0 L to pass through the DGM. | [10 empty] | Record the DGM
temperature and the absolute pressure of the DGM.
14.2.6 | Calculate the flow rate as follows:
|  |  |  | Q |  | /P | )(T | /T | )(F | )
|  |  |  | std | bar | std | std | m | c
where:
| Q std = | flow rate corrected to standard conditions, 25EC and 760 mm Hg, L/min.
| V = | volume of gas pulled through denuder system, L.
| t = | time required to pull gas through denuder system, minutes.
| Pbar = | barometric pressure, mm Hg.
| Pstd = | standard barometric pressure, 760 mm Hg.
| Tstd = | standard temperature, 298EK.
| Tm = | temperature of dry gas meter, EK(=EC + 273).
| F c = | dry gas meter correction factor, dimensionless.
14.2.7 | If the calculated flow rate is not between 9.5 and 10.5 L/min, readjust the flow rate and repeat
Sections 14.2.4 and 14.2.5 until the rate is in the above range. | [7 empty] | Preliminary studies should be conducted to
obtain an estimate of the concentrations of the species of interest.
14.2.8 | Record the flow rate on Field Test Data Sheet.
14.2.9 Remove DGM connection tubing from elutriator inlet. Pump should remain running so that sampling
continues. Higher flow rates may be used for shorter sampling periods. Concentration of the species of interest
in indoor air and the configuration of the sampling equipment, determine the appropriate flow rates. | [13 empty] | Sampling
at 10 L/min, requires a sampling time of 24 h to collect pollutant concentrations between 0.02 and 0.83 μg/m3 .
June 1999 |  |  | Compendium of Methods for Inorganic Air Pollutants | [10 empty] | Page 4.2-15
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
14.3 | Sample Shutdown
14.3.1 | Attach DGM connection tubing elutriator inlet with pump still running. |  |  | Measure flow rate as in
Sections 14.2.5 and 14.2.6. |  | Record flow time, temperature, and pressure on Field Test Data Sheet (see
Figure 13).
14.3.2 | Turn off pump. | Record time and elapsed time meter reading on log sheet. |  | Remove DGM
connection tubing from elutriator inlet. Remove ADS from the sampling box, cap the ends, and place the ADS
in field-to-lab carrying case for transport to lab. Be careful not to stress the ADS during the transfer or breakage
will result.
Caution: | When the ADS is brought from a cold field sampling location to a warm laboratory, loosen the
denuder couplings to prevent thermal expansion from breaking the denuders.
14.4 | Corrective Action for Leak Test Failure
[Note: These steps should be followed when failure occurs during testing at the laboratory before transport
to the field and in the field before testing.]
14.4.1 | Sampler Leaks. | Note the problem on the Field Test Data Sheet. |  | Check assembly of ADS
components. | Replace gaskets. | Check for proper seating of denuder surfaces. | Replace any defective parts.
14.4.2 | Cracked or Chipped Denuders or Elutriator Assemblies. |  | Note problem on Field Test Data
Sheet. | Discard defective pieces. | Do not try to extract cracked pieces. | WARNING--USE CAUTION WHEN
DISASSEMBLING CRACKED GLASSWARE. |  | Pieces may shatter and cause severe cuts. |  | Wear protective
clothing.
14.4.3 | Contaminated Blank Solutions. | Note problem on Field Test Data Sheet. |  | Follow parts-cleaning
procedures closely. | Examine the sampler preparation area for possible sources of contamination and remove
source. | Check DDW being used in the solution preparations and extractions. Fill a clean 25 mL polyethylene
extraction bottle with the DDW used in solution preparation and extraction; send to lab for analysis. |  |  |  | If
contaminated, correct deionization system.
14.4.4 Flow Rate Disagreement. |  | Note problem on Field Test Data Sheet. | Check vacuum gauge on flow
module. If a high vacuum exists, the sampler has become blocked. This blockage may be due to dust or smoke
particles clogging the filters or to obstructions in the system or tubing. |  |  | Check flow module. | Repair as needed.
14.4.5 | Inadequate Flow Rate. | Note problem on Field Test Data Sheet. |  | Check rotameter on flow
controller. | If adequate flow is shown here, a leak exists between the controller and the DGM. |  |  | If no flow is
shown on rotameter, check vacuum gauge on controller. |  |  | If no vacuum exists, pump needs repair. | If a high
vacuum is shown, an obstruction exists in the system. |  |  | Check to see that the paper filter dividers were not
accidentally installed with |  | the filters in the filter pack. | Check tubing for kinks.
[Note: Typically the pressure drop across the filters should be approximately 1" Hg at 10 L/min flow rate
at sea level. This pressure drop can vary from 1-10 L/min depending on elevation.]
15. ADS Disassembly
15.1 | Remove the ADS from the field-to-lab carrying case using both hands. |  | To prevent stress, hold the ADS
by its ends.
Caution: Do not stress the ADS while removing it from the case.
15.2 | Decouple the elutriator-jet assembly from the first denuder-impactor-coupler assembly.
Page 4.2-16 | Compendium of Methods for Inorganic Air Pollutants |  |  | June 1999
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
15.3 | When using the denuder-impactor, the frit-pin must be removed from the support in the denuder before
removing the frit from the pin. |  | The frit is then extracted from the pin using pin tool #3 and the frit extraction
tool (see Figure 12). When using the impactor-coupler assembly, the frit is removed from the coupler seat using
pin tool #3 and the "out" frit removal tool (see Figure 12). |  | Put frit in covered dish and set aside for future
chemical extraction (optional).
[Note: This method does not require the analysis of the particles impacted on the frit.]
15.4 | Remove the denuders from the couplers and cover each end of the denuders with clean end caps until
extraction.
15.5 | Label a clean 100 mL polyethylene bottle with the sampler ID number and filter type (i.e., Teflon® or
Nylasorb®, as appropriate) for each of the filters.
15.6 | Disassemble the filter pack in a clean, ammonia-free air hood. | Clean all hood surfaces and utensils with
ethanol. Wearing clean gloves and using clean filter forceps, remove the filters and place each in its storage
(protective) bottle, with the exposed filter surface facing downward, until extraction.
[Note: Place the filters in the properly labeled bottles.]
16. Extraction Procedures
[Special precaution: | Samples should be analyzed as soon after collection as possible. |  | The solutions and
extraction procedures must be prepared and performed on the day of pH analysis.
Extraction must take place in a clean, ammonia-free air hood. The extracts must be processed in the order
in which they will be analyzed, so that each sample will have a similar time interval between extraction and
analysis. Denuder extracts and filters should be stored in the refrigerator until just prior to analysis. Samples
stored longer than 30 days tend to degrade due to bacteria growth and/or losses to the walls of the extraction
vessel. | Extraction volumes may vary depending upon sensitivity requirements and expected ambient
concentrations.
16.1 | Impactor Frit Coating Extraction (Optional)
16.1.1 | Place the impactor (which was removed before denuder extraction) into a small extraction bottle.
16.1.2 | Label the bottle appropriately. | Pipet 10 mL of impactor extraction solution (See Section 9.2) into
the bottle. | The solution must cover the surface of the impactor frit.
16.1.3 | Close the extraction bottle and place in an ultrasonic bath for 30 min.
16.2 | Denuder Extractions
[Note: If the denuder was the first denuder and equipped with the impactor frit-pin support, insert a clean
Teflon® impactor frit-pin without frit in place. |  | Then extract as described below. | This procedure is to be
followed for each denuder.]
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-17
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
16.2.1 | Cap one end of each of the denuders (i.e., Na 2 CO3 |  | and citric acid denuders). | Add 5 mL of DDW
with a pipet. | Cap other ends of the denuders.
16.2.2 | Rotate the denuders to wet all surfaces thoroughly with the water. |  |  | Remove the caps and pour the
liquid from each of the two denuders into separate clean 25 mL polyethylene extraction bottles.
16.2.3 | Repeat this procedure with a second 5 mL of DDW extract (total extract volume is 10 mL, which
is placed into the extraction bottles).
16.2.4 Add 1 mL of hydrogen peroxide to oxidize the sulfite (SOG3 ) to sulfate (SO=4) in the extraction bottles
containing the Na2 CO3 |  | coating.
16.2.5 | Replace the extraction bottle caps and label the bottles with the sampler ID number and denuder
number and type (as appropriate).
16.3 | Filter Extraction
16.3.1 | Teflon® Filter Extraction [for pH analysis followed by SO=4, NOGG3 , and/or NH +4 analyses)
[Note: Teflon® is not wet by water; therefore, the filter will float on top of aqueous solutions. Solutions and
extraction procedures must be prepared and performed on the day of pH analysis. Extraction of the filters
must take place in a clean, ammonia-free, air hood. The filters must be processed in the order in which they
will be analyzed so that each sample will have a similar time interval between extraction and analysis.]
| 16.3.1.1 | Samples should be analyzed as soon after collection as possible. The solutions and extraction
procedures must be prepared and performed on the day of pH analysis. |  |  |  | Keep samples in a refrigerator until
extracted and analyzed.
| 16.3.1.2 | Samples should not be extracted until the day of analysis; however, if samples are extracted
and it is not possible to analyze them that day, they should be refrigerated. Allow the samples to return to room
temperature before analysis.
| 16.3.1.3 | The same extract solution (ES) must be used for the samples to be |  | analyzed, the working
standards, and the EA solution. |  |  | Also, the same batch of alcohol must be used to prepare the EA solution and
the working standards and to extract the Teflon® filters.
| 16.3.1.4 | Handling the exposed Teflon® filter requires protection from contamination with NH3 , which
may rapidly neutralize aerosol acidity on the filter and bias the sample results. |  |  |  | To ensure ammonia-free air
occupies the glove-box, a positive pressure is maintained by blowing air through a PVC tube (4" O.D.) filled
with glass-wool dosed with citric acid before entering the manifold which enables uniform distribution of air
from top of the glove-box. |  | Flow the ammonia-free air for 5 minutes before retrieving the filter. |  | Place a citric
acid soaked filter paper on the bottom to deplete ammonia when unused. |  |  |  | Disassemble the filter assembly in
the clean, ammonia-free glove-box. |  |  | Clean all glove-box surfaces and utensils with ethanol.
| 16.3.1.5 | Allow the hood to be flushed with ammonia-free air for at least 5 min before filter extraction.
All of the hood surfaces and extraction utensils must be cleaned with a Kimwipe® moistened with ethanol.
| 16.3.1.6 | Pipet 6 mL of 0.0001 N perchloric acid (HClO 4 ) solution into the appropriately labeled
extraction vials (10 mL).
[Note: Perchloric acid is used because it inhibits CO 2 from dissolving into the solution and keeps the organic
compounds in solution from dissociating. Both these activities can change the ionic strength of solution.]
| 16.3.1.7 | Wearing clean gloves and using clean filter forceps, place the Teflon® filter in the extraction
vial with the exposed filter surface facing downward. |  |  | Cap tightly. | Store at 4EC in the dark until ready for
analysis.
Page 4.2-18 |  | Compendium of Methods for Inorganic Air Pollutants |  | June 1999
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
16.3.1.8 | When ready for analysis, the filter must be prepared (within the air hood) in the following
manner: | Using forceps and gloved hands, lift the filter from the extraction vial. | Let the excess solution drain
off into the vial. | Holding the filter over the extraction vial, and using an automatic pipet, apply 200 FL ± 5 FL
of ethanol to the filter. Add the ethanol slowly to ensure that all portions of the membrane are wet with ethanol.
Immerse the filter in the aqueous solution once again. |  | Tap the forceps against the inside of the vial to remove
liquid. | Tightly replace cap. | Put in ultrasonic bath for 15 min total, rotating the rack 90E every 5 min.
[Note: | Perchloric acid is used in place of potassium chloride, initially, to prevent interference in the
measurements of cations and anions by ion chromatography. |  | Potassium chloride must be added to the
portions of the sample extract that are used for pH analysis (the purpose of the salt, final concentration
0.04 M, is to increase the ionic strength and thus to reduce the time for equilibrium of the pH electrode used
for measurement). Use the same bottle (freshly opened) of ethanol to extract the Teflon® filters that are used
to prepare sulfuric acid standards.]
16.3.1.9 When ready for pH analysis, the extracts are prepared in the order of pH measurement. Inside
the air hood, remove the caps from 10 mL extraction vials. |  | Wipe off any drops which may leak onto the
outside of the cup.
16.3.1.10 | Using gloved hands and a 1 mL automatic pipet, transfer 1 mL of the extract to each of two
correspondingly labeled 2 mL cups.
[Note: The first 2 mL cup for each extract has the same I.D.# as the 10 mL cup, and the second 2 mL cup
has the same I.D.# with a hyphen (-). This system also is used with the working standards.]
16.3.1.11 | After transferring the extracts to the 2 mL cups, recap the 10 mL extraction vial. | Store the
10 mL vials at 4EC in a refrigerator pending analysis by IC, and for NH4+ analysis by Technicon auto analyzer.
16.3.2 | Nylon Filter Extraction.
16.3.2.1 | Pipet 10 mL of IC eluent into the appropriately labeled filter vial or bottle with caps.
[Note: Be sure that the filter lies flat on the bottom of the bottle and that all of the filter is covered by the
extraction solution.]
16.3.2.2 | Place the nylon filter in the extraction vial.
16.3.2.3 | Replace the bottle's cap and put in an ultrasonic bath for 30 min.
16.3.2.4 | Store the bottles in a clean (i.e., pollutant free) refrigerator at 4EC until analysis.
17. Ion Chromatography Analysis
[Note: | The analytical procedure described here is | not the only appropriate procedure available for
quantifying the analytes of interest. An automated system does not have to be used. This particular analytical
procedure was chosen because it is presently used by EPA. Modifications to this procedure may be required
depending on the intended use of the data; however, any modifications made must be justified to obtain
comparable data quality.]
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-19
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
17.1 | Standards Preparation
Special Precaution: |  | Storage of these solutions should be no longer than one week. |  | All of the working
standard solutions are used to calibrate the IC and are made from reagent grade stock. The crystals are dried
overnight in covered petri dishes at 110EC in a vacuum oven prior to preparing the standard solutions. Any
yellowish discoloration of the dried crystals indicates decomposition and crystals should be discarded.
17.1.1 | Sodium Sulfate Stock Solution.
| 17.1.1.1 | In a clean, calibrated, 1 L flask, add 500 mL of DDW.
| 17.1.1.2 | On weighing paper, weigh out enough reagent (Na2 SO4 ) to make the solution 2000 ppm
concentration. |  | The target weight is 2.958 ± 0.001 g. | Record the gross weight.
17.1.1.3 Add the reagent crystals to the 500 mL of DDW. Reweigh weighing paper and subtract weight
from the gross weight. |  | The difference is the actual net weight.
| 17.1.1.4 Dilute to volume with DDW. | Mix well and cover with parafilm.
17.1.2 | Sodium Nitrate Stock Solution.
| 17.1.2.1 | In a clean, calibrated, 1 L flask, add 500 mL of DDW.
| 17.1.2.2 | On weighing paper, weigh out enough reagent (NaNO3 ) to make the solution 2000 ppm
concentration. |  | The target weight is 2.742 ± 0.001 g. | Record the gross weight.
| 17.1.2.3 | Follow Sections 17.1.1.3 and 17.1.1.4.
17.1.3 | Standard working solutions (100 ppm SO 4= and 100 ppm NO3- ). |  |  | The working solutions are
made as follows: Add 10 mL each of the two stock solutions (SO4= and NO 3 G) to a 200 mL volumetric flask
and dilute to the mark with DDW. |  | Subsequent dilutions are carried out using a 10 mL volumetric pipet and
appropriate flasks. |  | Standards of 20, 10, 5 and 1 ppm | SO4= and NO 3 G are prepared to calibrate the IC.
17.2 | Reagent Preparation
[Note: Storage of these reagents should be no longer than 1 week.]
17.2.1 | Anion eluent. | The anion eluent is a solution of 1.8 mM Na 2 CO3 |  | and 1.7 mM NaHCO 3 . | A
concentrated solution can be prepared and diluted as needed.
[Note: See Anion Storage Solution]
| 17.2.1.1 Concentrated Na 2 CO 3 | solution (0.36 M). | Weigh out 38.156 g of Na2 CO3 (MW = 105.99).
Dissolve into 1 L of DDW. |  | Store in refrigerator until ready to dilute.
| 17.2.1.2 Concentrated NaHCO 3 solution (0.34 M). |  | Weigh out 28.564 g of NaHCO3 (MW = 84.01).
Dissolve into 1 L of DDW. |  | Store in refrigerator until ready to dilute.
| 17.2.1.3 Dilution of stock solutions. | Bring both solutions to room temperature. |  | Accurately pipet 10
mL of each solution into a 2000 mL volumetric flask which has been partially filled with DDW. |  |  |  | Bring to the
mark with DDW (1:200 dilution).
17.2.2 | Anion regenerant. | The regenerant is a 0.025 N H 2 SO4 solution. | VERY CAREFULLY dispense
2.8 mL of concentrated Ultrex sulfuric acid (36 N) into a graduated cylinder. |  |  |  | Partially fill the regenerant
reservoir with DDW (3 L). |  | Slowly add the acid to the regenerant reservoir. | Bring to the mark with DDW
(4 L).
[Note: Protective clothing and eye protection should be worn.]
Page 4.2-20 |  | Compendium of Methods for Inorganic Air Pollutants |  | June 1999
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
17.2.3 | Cation eluent. | There are two cation eluents that are used for the analysis of monovalent and
divalent cations. |  | The strong cation eluent is: 48 mM HCl, 4 mM DAP.HCl (DAP = Diaminoproprionic acid),
4 mM Histidine.HCl. |  | The weak eluent consists of 12 mM HCl, 0.25 mM DAP.HCl, 0.25 mM Histidine.HCl.
| 17.2.3.1 Strong cation eluent. |  | Weigh 0.560 g DAP and 0.840 g histidine into a 1 L volumetric flask.
Add 48 mL of 1 M HCl (Ultrex) to the flask. |  |  | Bring the eluent to the final volume by bringing to the mark with
DDW. | Mix thoroughly to dissolve.
| 17.2.3.2 | Weak cation eluent. | Place 63 mL of the strong cation eluent in a 1 L flask. Add 9 mL of 1
M HCl to the flask. |  | Bring the eluent to the final volume by bringing to the mark with DDW. Mix thoroughly
to dissolve.
17.2.4 | Cation regenerant. | The cation regenerant consists of 100 mM tetrabutyl-ammonium hydroxide
(TBAOH). Place the TBAOH container into a warm water bath to dissolve any crystals that may have formed.
Measure 266.7 mL of the TBAOH (stock reagent is supplied as 1.5 M, 40% in water) into a graduated cylinder.
Add the TBAOH to 4 L of DDW.
17.2.5 | Anion storage solution. | Since the anion columns contain carbonates from the eluent, protection
must be taken against microorganisms that will live on this food source and clog up the columns. If the columns
are not being used for long periods of time (>2 weeks), a storage solution of 0.1 M NaOH should be pumped
into them.
17.3 | Sample Preparation
17.3.1 | Mark the auto sampler vials with the appropriate identification numbers. |  | Place the vials in an IC
autosampler tray.
17.3.2 | Using clean, calibrated 0.5 mL pipets, transfer the denuder extract and the remainder of the filter
extracts from the extraction vials to a clean disposable 0.5 mL IC autosampler (polyethylene) vial. |  |  | Fill the
autosampler vial to the line on the side.
[Note: If refrigerated, the contents of the 10 mL extraction vial must be vortex-mixed prior to transfer to the
autosampler vials.]
17.3.3 | Place black filter caps on top of the vials. |  | Use the tool provided to push the caps into the vials until
they are flush with the top |  | (see the IC manual for more detailed instructions).
17.3.4 | Wipe away any excess fluid from the top of the vial to avoid contamination from other samples.
17.3.5 | After all of the trays are filled, place them into the left side of the autosampler. |  | The white dot on
the tray indicates the first sample. Press the button labeled RUN/HOLD to the RUN position. The trays should
move until the first sample is under the sampling head. |  |  | The front panel should indicate a READY message.
Press local/remove switch to remove.
17.4 | Basic System Operations - Start-up and Shut-down
[Note: This procedure is specific for one instrument. Other instruments can be used with these procedures.]
17.4.1 | Start-up Procedure for Ion Chromatograph.
| 17.4.1.1 | The major components of the Dionex 2020i Ion Chromatography system are illustrated in
Figure 14. Turn helium and nitrogen tanks on by opening the valve on top of each tank (pressure in either tank
should not be less than 500 psi.). |  | Replace if necessary. | Open valves at the outlet end of both regulators.
Adjust pressure on the nitrogen regulator to 100 psi. |  |  | Adjust pressure on the helium regulator to 14 psi.
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-21
Method IO-4.2 |  |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  |  | Atmospheric Acidic
17.4.1.2 Check the level of eluents and regenerating solutions. Turn the chromatography (CMA) values
for the anion channel switch ON. |  |  | Verify that the pressure reading on the face of the degassing unit is 7 psi.
Adjust by turning dial next to pressure gauge. |  |  |  | Turn the degas switch to HIGH.
| 17.4.1.3 |  | Turn the eluent reservoir switches, corresponding to the eluents to be degassed, to the ON
position. |  | Let the eluents degas on HIGH for 3-5 min, then turn degas switch to LOW.
| 17.4.1.4 |  | Select the appropriate program on the gradient pump module using the PROGRAM switch.
(Programs are recalled from memory by first pressing the PROGRAM switch, then the single digit reference
number corresponding to the appropriate program.)
| 17.4.1.5 |  | Prime the eluent lines.
[Note: | All of the eluent lines used during analysis must be primed to remove any air bubbles that may be
present. The selected program identifies which lines are used.]
| • | Open the gradient pump drawer. Turn the pump to the START position for 10 s, or until a CLICK
is heard. Turn the pump OFF. This step opens the valve to the eluent line displayed on the front
panel.
| • | Attach a 10 mL syringe to the priming block on the face of the gradient pump module. With the
priming block valve closed, pull the syringe plunger out to the end of the syringe.
| • | Open the priming block valve. |  | The syringe will quickly fill with eluent. | Close the valve on the
priming block when the syringe is almost full. Remove syringe from block and discard collected
eluant.
This priming procedure can be repeated if necessary. All of the eluent lines that are to be used during a day
of analysis should be primed at this time.
| 17.4.1.6 |  | Open the door of the Advanced Chromatography Module. | On the back of the door, at the
bottom, is the conductivity detector. |  |  | Four labeled lines (anion, cation, waste, and cell) are located next to the
cell. | The plumbing must be configured according to the type of analysis to be performed. |  |  |  | If anions are being
analyzed, attach the ANION line to the CELL line, and the CATION line to the WASTE line. |  |  |  |  | If cations are
being analyzed, attach the CATION line to the CELL line, and the ANION line to the WASTE line. Attach the
line coming from the pump to the correct port on the advanced chromatography module. |  |  |  |  | SYSTEM 1 on the
left is for anions; SYSTEM 2 on the right is for cations.
[Note: | If switching from one system to the other, the pump and the lines coming from the pump must be
purged of the original eluent, which is done by disconnecting the pump line from the chromatograph module,
turning the pump on, and running the new eluent into a waste beaker for 2-3 min.]
| 17.4.1.7 |  | Select the columns to be used (labeled pH or NO2 ) by pressing the blue button located below
the labels. |  | To verify that the correct columns are being used, press the switch should at least once, and then
set to the appropriate position.
| 17.4.1.8 |  | Turn the power switch on the autosampler ON (switch is located on the back of the unit, on
the right). |  | The default settings will be displayed on the front panel. |  | Attach the SAMPLE OUT line from the
autosampler to the advanced chromatography module. |  |  |  | The connection should be made to the port marked
SAMPLE of the appropriate system. |  |  | Turn the pump to START.
| 17.4.1.9 |  | Turn the conductivity cell ON (switch is located on the gradient pump module). |  | Turn the
REGEN switch for the appropriate system ON. |  |  |  | Verify that regenerant is flowing by inspecting the regenerant
waste line which empties into the sink. Open the advanced chromatography module door and inspect for leaks
at columns, fittings, etc. |  |  | Shut pump off if leaks are found.
Page 4.2-22 |  |  | Compendium of Methods for Inorganic Air Pollutants |  | June 1999
Chapter IO-4 |  |  |  |  | Method IO-4.2
Atmospheric Acidic |  |  |  | Acidic/Basic Constituents
| 17.4.1.10 | Turn stripchart recorder ON. | Baseline should stabilize in less than 20 min. If baseline is not
stable, see troubleshooting Section 17.5 for assistance.
17.4.2 | Data acquisition start-up. | The following is a description of the current data acquisition program
used by the EPA. |  | The program is available (EPA, Atmospheric Chemistry and Physics Division, Office of
Research and Development, Research Triangle Park, NC) and is for IBM or IBM compatible computers. Other
appropriately designed programs may be used to compile the data collected for any given sampling network.
A computer programmed integrator does not need to be used to compute the data, but is recommended for large
sampling networks.
17.4.2.1 Turn on the IBM XT computer. From the C:>prompt, type: cd/cchart, then type: cchart. This
loads the Chromatochart software. |  | Turn switch on relay box to ENABLE, indicator light could go on.
| 17.4.2.2 Press F2 to enter the methods development module. |  | Select option number 1 - "select channel
# and load method file." |  | "Select channel # " type 0 or press ENTER to select the default choice shown in
the brackets (in this case 0). |  | "Load method file named" type the name of the appropriate method, then press
ENTER. | A directory of all of the current methods in memory can be obtained by pressing the F2 function key.
| 17.4.2.3 | Press F3 to enter the Data Acquisition module. | At this point you will be asked to save the
method file. | If there has not been any changes to the methods file, it does not need to be saved. Select option
#4 - "Collect Data." Press ENTER to deactivate the method queue. "Load Run Queue named," type the name
of the run queue if one has been created. |  | Type ENTER to deactivate the run queue.
| 17.4.2.4 | "Total # runs for method ," type how many times the method is to be repeated (total
number of samples). |  | "Autoanalyze Data" type Y. | "Autosave data to | disc" type Y. | "Data file name (xxxxx)
change?", type data file name. |  | "Press ENTER to begin methods." | Press ENTER only after the samples have
been loaded into the autosampler and the baseline has stabilized.
17.4.2.5 Figure 15 illustrates the chromatograms for each of the samples as output by the programmed
Spectra-Physics integrator.
17.4.3 | Calibration of IC. | The instrument should be brought to normal conditions with a warm-up time
of at least 30 min.
17.4.3.1 With the "Reading" light on, check to ensure the flow rate is 1.5 mL/minute, the fluid pressure
is 600 psi ± 100 psi and the conductivity is constant as measured by offset difference.
17.4.3.2 Fill the IC vials with the prepared standard solutions and (20, 10, 5 and 1 ppm SO4= and NO 3 G)
and pure eluent. |  | This will allow a four-point calibration curve to be made.
[Note: For low-level applications, more standards and blanks may be necessary in order to obtain accurate
reference curves.]
| 17.4.3.3 | Load the four vials into the sample vial holder and place the holder in the automated sampler
tray.
| 17.4.3.4 The tray is controlled by a Spectra-Physics SP4200 or SP4270 Computer Integrator. |  |  |  | Use the
integrators operation manual to begin calibrating. By using the RUN command, the analysis and data treatment
phases of the calibration are set in motion. |  | Four calibration standards are run, the chromatograms and peak
areas displayed for each run, and the run results for each anion are fitted to a quadratic curve by a least squares
regression calculation. The three curves are plotted and the correlation coefficients are calculated. The values
of the coefficients are normally greater than 0.999, where 1.000 indicates a perfect fit. Values of less than 0.99
indicate the calibration procedure should be repeated.
[Note: Recalibration should be carried out whenever standard concentrations show consistently high or low
results relative to the calibration curve is compared to the calibration curve from the old standards.
Comparability of points should be within ± 0.1 ppm or ± 10%. For standard concentrations of greater than
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants |  |  | Page 4.2-23
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
1 ppm, comparability will normally be within 5% or better. Old standards are assumed correct since they
are referenced to the entire historical series of previous standard solutions all of which are comparable.]
17.4.4 |  | System Shut-down.
| 17.4.4.1 | Shut off the pump. | Turn the REGEN switch and the conductivity cell to the OFF position.
| 17.4.4.2 | Switch the eluent degas switch to HIGH.
| 17.4.4.3 Turn the stripchart recorder OFF. |  | Cap the pen. | Press the F10 function key on the computer.
Select option 3 to exit to DOS. |  |  | Shut off the printer and the computer.
17.4.4.4 Shut the eluent degas system and reservoir switches and the autosampler to the OFF position.
Close the valves on both gas cylinders. |  |  | Close the regulator valves.
17.5 | Basic Troubleshooting
Before proceeding with the troubleshooting guide, make sure the reagents used were prepared correctly and
are not "old."
17.5.1 |  | Unstable Baselines
| 17.5.1.1 Wavy baseline. |  | The most common reason for a wavy baseline is an air bubble in the gradient
pump. | This problem is diagnosed by observing the pump head indicator lights on the gradient pump module
front panel. |  | If the baseline is pulsing in phases with pump pistons, a bubble is usually present. |  | Other
possibilities include a dirty or stuck check valve, piston seal or "O"-ring, as well as an air bubble in the
conductivity cell.
17.5.1.2 Drifting baseline. Steadily increasing or decreasing baselines usually indicate that the suppressor
column is not performing as it should. |  |  | Parameters to change include the regenerant and eluent concentrations
and flow rates. |  | Check temperature routinely, as changes in temperature can cause drifting. |  | Balancing these
should stabilize the baseline if the suppressor is functioning correctly. |  |  |  | The Dionex manual describes clean-up
procedures if the suppressor is believed to be contaminated.
| 17.5.1.3 | High baselines. | As with drifting baselines, the parameters to change are eluent and regenerant
concentrations and flow rates. |  |  | A high baseline usually indicates that there is not enough baseline suppression;
this condition can be controlled by increasing the regenerant flow rate.
| 17.5.1.4 | Low baselines. | Low baselines usually indicate that there is too much suppression, which can
be controlled by decreasing the flow of the regenerant.
17.5.2 |  | Backpressure. | Variations in system backpressure are common and should not raise concern
UNLESS the pressure change is greater than 200 psi.
| 17.5.2.1 | High backpressure. | The system is protected from pressure related damage through the high
and low pressure alarm settings on the front panel of the gradient pump module. |  |  |  | If the high pressure setting
is correctly selected (200 psi above normal operating range), the pump will automatically shut-off if this value
is exceeded. |  | The reason for high backpressure is blockage in the system. |  | Possibilities include loading against
a closed valve, a plugged line, contaminated columns, etc. |  |  |  | Diagnosis is done by removing one component of
the system and observing how the pressure changes.
| 17.5.2.2 | Low pressure. | Low pressure readings usually indicate a leak somewhere in the system.
Carefully check all fittings for leaks. |  |  | Tighten if necessary.
17.5.3 |  | Flow
| 17.5.3.1 | Regenerant lines. | If there is no flow at the waste outlet end of the regenerant line, check the
following:
| • | Make sure the correct regenerant switch is turned on
| • | Verify that the reservoir is not empty
| • | Make sure the nitrogen tank is turned on
| • | Check that the regulator is correctly set
Page 4.2-24 |  |  | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 | [7 empty] | Method IO-4.2
Atmospheric Acidic | [5 empty] | Acidic/Basic Constituents
| 17.5.3.2 | Eluent lines. | If there is no flow at the outlet end of the eluent lines, check the following:
| • | Check that the pump is on
| • | Check that the eluent lines are connected to the correct port
17.5.4 |  | Software. | Refer to the ChromatoChart manual for detailed information on software problems.
18. Ammonia Analysis by Technicon Autoanalysis
Presented in Sections 18.1 and 18.2 are the recipes for the standards and reagents required to analyze the
ammonium ion (NH4+ or ammonia (NH 3)) by Technicon autoanalysis. The prelude of these sections briefly
describes the TRAACS 800 autoanalyzer and the sample flow through the TRAACS 800 for NH 4+ analysis.
This instrument is capable of quantifying, from a single sample, three different species simultaneously. An
aliquot of the sample is taken from an automated sampler by syringe. A splitter divides the aliquot into the
appropriate volumes required for the particular analyses. |  |  | Each of the volumes is then transferred to the
appropriate analytical cartridge. |  |  | Sample flow diagrams that illustrate SO 4= , NO3- and NH4+ analysis
(although in Method IO-4, only ammonia is analyzed by this technique) from the citric acid denuder and
Teflon® filter can be shown separately and independently of one another. | [5 empty] | The data computation (by
computer) and quality assurance protocols, however, can not be readily adapted to single-channel
instruments. |  | These protocols need to be specific to the individual analytical instrument. | [5 empty] | In brief, NH 4+
analysis is illustrated in Figure 16. The samples, along with all standards, are taken from the auto-advance
sampler tray by the use of a proportioning pump and automated syringe. Air and EDTA are first added to
the samples and are mixed in the first set of coils. After mixing, phenolate is added and mixed in the next
set of coils. Nitroprusside is then added and mixed, followed by the addition and mixing of hypochlorite.
At this stage, the sample should be a bright blue color. After the last mixing stage, the sample is sent through
a heated bath, followed by another mixing stage. Finally, the sample is sent through a colorimeter where the
results are recorded on a digital printer and stored in a computer file for further manipulation.
[Note: Ammonia can be analyzed by other techniques.]
18.1 | Standards and Stock Solutions Preparation
[Note: | Before discarding the old solution, it should be checked against the fresh solution by comparing
calibration curves on the working solutions prepared from them. |  |  |  | Slopes and intercepts are calculated for
each set of standards. The old slope and intercept are used to calculate concentration values from readings
for the new standards. This calculation determines if the old solution has deteriorated or if an error has been
made in preparing the new solution.]
18.1.1 |  | Ammonium Solution Standard (1,000 μg/mL). | Dry ammonium chloride in an oven for 1 h at
50-60EC and desiccate over silica gel for 1 h. |  |  | Weigh 2.9470 g ammonium chloride and dissolve in 800 mL
DDW. | Dilute to 1 L with DDW and mix thoroughly. |  | This solution is stable for 1 y.
18.1.2 |  | Intermediate Ammonium Standards. | To make a 100 μg/mL ammonium standard, pipet 10 mL
of ammonium stock standard into a 100 mL volumetric flask. Dilute to volume with DDW and mix thoroughly.
Keep refrigerated. This solution remains stable for 1 month. To make a 10 μg/mL ammonium standard, pipet
1.0 mL of ammonium stock standard into a 100 mL volumetric flask. |  |  |  | Dilute to volume with DDW and mix
thoroughly. |  | This solution remains stable for one week.
18.1.3 |  | Working Ammonium Standards in DDW. | Pipet | aliquots | of | the | 100 | μg/mL | ammonium
intermediate standards into 100 mL volumetric flasks according to the table below. | [6 empty] | Dilute to volume with
DDW. | Prepare fresh daily.
June 1999 |  |  | Compendium of Methods for Inorganic Air Pollutants |  |  |  |  | Page 4.2-25
Method IO-4.2 |  |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
|  | Stock or intermediate |  | Concentration,
| Standard | standard, Fg/mL | Aliquot, mL | Fg/mL
| A | 1,000 | 4.0 | 40.0
| B | 100 | 4.0 | 4.0
| C | 100 | 3.0 | 3.0
| D | 100 | 2.0 | 2.0
| E | 100 | 1.0 | 1.0
| F | 100 | 0.5 | 0.5
| G | 10 | 2.0 | 0.2
| H | 10 | 1.0 | 0.1
18.2 | Reagent Preparation
[Note: When reagents are prepared, label the container with the contents, concentration, date prepared, and
the preparer's initials.]
18.2.1 | Alkaline Phenol. | Add 83.0 g loose crystallized phenol to 800 mL DDW in a 1 L volumetric flask.
Keeping the flask in an ice bath or under tap water, slowly add 96.0 mL 50% sodium hydroxide solution. Shake
the flask while adding the sodium hydroxide. |  | Cool to room temperature, dilute to 1 L with DDW, and mix
thoroughly. | Store in an amber glass container. | This solution will remain stable for 3 mo, if kept out of direct
light.
18.2.2 | Sodium Hypochlorite Solution. | The amount of sodium hypochlorite solution varies from batch to
batch of sodium hypochlorite (5% commercial grade). |  |  | Therefore, for each new batch, a base and gain
experiment must be run to adjust the amount of sodium hypochlorite required to obtain the existing base and
gain values. | In a 150 mL volumetric flask, dilute 86 mL of 5% sodium hypochlorite solution to 100 mL with
DDW and mix thoroughly. Check base and gain values. Reduce or increase the amount of sodium hypochlorite
to obtain the same base and gain values as the previous sodium hypochlorite batch. This solution remains stable
for 1 day.
18.2.3 | Sodium Nitroprusside Solution. | Dissolve 1.1 g of sodium nitroprusside in about 600 mL of DDW,
dilute to 1 L with DDW, and mix thoroughly. |  | Store in an amber container and keep in refrigerator. |  |  | This
solution remains stable for 1 month, if kept out of direct light.
18.2.4 Disodium EDTA Solution. Dissolve 1.0 mL of 50% w/w sodium hydroxide and 41.0 g of disodium
EDTA mix thoroughly. | Add 3.0 mL of Brij-35 and mix. |  | Store in plastic container. | This solution will remain
stable for 6 months.
19. pH Analysis
19.1 | Standard and Reagent Preparation
[Note: Each of the standard H 2SO 4 stock solutions must be prepared fresh the day of pH analysis.]
Page 4.2-26 | Compendium of Methods for Inorganic Air Pollutants |  |  |  | June 1999
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
19.1.1 | Standard H2 SO4 | Solution.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-27
Method IO-4.2 | [5 empty] | Chapter IO-4
Acidic/Basic Constituents | [5 empty] | Atmospheric Acidic
| 19.1.1.1 | Label seven 25 mL polyethylene stoppered volumetric flasks. |  |  | Also, label each flask with the
volume of 1 N H2 SO4 |  | solution indicated in the following table:
Volume of
|  |  | 1.000 N H2SO 4 |  | Standard
|  | Standard H 2SO 4, | added to each |  | concentration,
|  | Flask No. | flask, μL |  | 10 -3N H2SO 4
|  | 1 |  | 0 | 0
|  | 2 |  | 25 | 1
|  | 3 |  | 50 | 2
|  | 4 |  | 100 | 4
|  | 5 |  | 200 | 8
|  | 6 |  | 400 | 16
|  | 7 |  | 800 | 32
| 19.1.1.2 | Use the 25 μL automatic pipet to add 1 N stock H2 SO4 |  |  | to flasks #1-3. | Use the 100 μL pipet
to add 1 N stock H 2 SO4 |  | to flasks #4-7. | Dilute all flasks to the 25 mL mark with absolute ethanol. |  |  | Cap with
stoppers or parafilm and mix well.
19.1.2 | 2 M Potassium Chloride (KCl) Solution.
| 19.1.2.1 | Weigh 149.2 ± 0.1 g of KCl. | Add the KCl to a 2 L flask.
| 19.1.2.2 | Add about 700 mL of DDW water to the flask. |  |  | Swirl the solution until the KCl is completely
dissolved.
| 19.1.2.3 | Pour this mixture into a 1 L graduated cylinder. |  |  | Rinse the flask with a small amount of water
and transfer the rinse into the cylinder. |  |  | Fill the cylinder to the 1 L mark.
| 19.1.2.4 Pour the solution from the cylinder into the 1 L polyethylene bottle. |  |  |  |  | Cap and shake the bottle
to mix well. | Mark the bottle with date of preparation.
19.1.3 | 0.1 N Perchloric Acid (HClO4 ) Solution.
19.1.3.1 Fill a 1 L graduated cylinder about 1⁄2 full with DDW. Transfer 10 ± 0.1 mL of 60-62% HC104
into the 1 L cylinder with a 10 mL pipet.
| 19.1.3.2 | Fill the cylinder to the 1 L mark. | Pour the solution into the 1 L polyethylene bottle.
| 19.1.3.3 | Cap and shake the bottle to mix well. |  | Mark the date of preparation on the bottle.
19.1.4 | 0.01 N HClO 4 | Solution.
| 19.1.4.1 | Fill a 1 L graduated cylinder about 1⁄2 full with DDW.
| 19.1.4.2 | Measure 100 mL of the 0.1 N HC1O 4 |  |  | solution with the 100 mL graduated cylinder. | Add
solution to the 1 L cylinder.
| 19.1.4.3 | Fill a 1 L cylinder with DDW to the 1 L mark. |  |  | Pour the solution into the 1 L polyethylene
bottle.
| 19.1.4.4 | Cap and shake the bottle to mix well. |  | Mark the date of preparation on the bottle.
19.1.5 | Extraction Solution (0.0001 N HClO 4 ).
[Note: This solution has the same composition as the solution used to fill the sample vials for Teflon® filters.
It must be prepared fresh on the day of pH analysis.]
| 19.1.5.1 | Using a 5 mL calibrated automatic pipet, add 10 ± 0.1 mL of 0.01 N perchloric acid (HClO 4 )
to a 2L Erlenmeyer flask of water. |  |  | Add 990 ± 10 | mL of DDW to the flask.
Page 4.2-28 |  | Compendium of Methods for Inorganic Air Pollutants |  |  |  | June 1999
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
| 19.1.5.2 | Mix well and cover with parafilm until ready for use.
19.1.6 | EA Solution.
19.1.6.1 Measure 150 ± 2 mL of 0.0001 N HClO 4 (prepared in Section 19.1.5) into a 250 mL graduated
cylinder. | Transfer to a 250 mL Erlenmeyer flask.
| 19.1.6.2 | Using a 5 mL graduated cylinder, add 5 ± 0.1 mL of ethanol (from the same fresh bottle of
ethanol that was used to prepare the standards in 19.1.1) to the flask.
| 19.1.6.3 | Again using a 5 mL graduated cylinder, add 3 ± 0.1 mL of 2 M potassium chloride (KCl)
solution (see 19.1.2) to the flask.
| 19.1.6.4 | Mix well and cover with parafilm until ready for use.
19.1.7 | Working Standard Test Solutions.
| 19.1.7.1 | Place fourteen-4 mL polystyrene sample cups (as used with Technicon Auto-Analyzer II
system) labeled 1, 1*, 2, 2*...7, 7* into racks. |  | Using the calibrated dispensing pipet bottle, add 3 mL of EA
solution to each 4 mL cup.
| 19.1.7.2 | Using the displacement pipet, add 50 μL of absolute ethanol to each cup. | Pour about 3 mL
of standard (H 2 SO4 |  | solution) #1 into a labeled 4 mL cup.
| 19.1.7.3 | Immediately pipet 50 μL of this standard into the 4 mL cups labeled 1 and 1* containing the
EA solution and ethanol.
[Note: | This transfer must be done without delay to prevent the standard concentration from increasing
significantly due to evaporation of the ethanol solvent.]
| 19.1.7.4 | Repeat the procedure for each of the other 6 standards. If there is a delay of more than 5 min
between the preparation of these mixtures and the next step, put caps on the 4 mL cups.
| 19.1.7.5 To prepare for analysis, mix each mixture. |  | Then transfer two aliquots from each cup to 2 mL
sample cups. | Place cup #1 in a rack. | In a second rack place two-2 mL cups labeled 1 and 1-. | Use the 1 mL
automatic pipet to mix the contents of 4 mL cup #1 by drawing 1 mL into the pipet tip and then dispensing it
back into the 4 mL cup three times. |  | Use the same pipet to transfer 1 mL of the mixture to each of the two
labeled 2 mL cups. |  | Place caps on the two 2 mL cups. | After transferring the two aliquots to 2 mL cups, rinse
the automatic pipet tip in a flask of DDW. Repeat the transfer procedure for each of the other working standard
pairs.
19.2 | Calibration of pH Meter
The pH meter requires temperature calibration whenever a new electrode is used. |  |  | Use the manufacture's
procedure in the instrument manual. This calibration should be repeated every 3 months when not in use.
The pH meter is left with the power cord plugged into the AC outlet, the mode control knob is left in the
standby position, the electrode lead is partially disconnected by pressing the plastic ring on its outer edge,
and the combination electrode is immersed in a 4 M KCl solution (a slit rubber stopper seals the bottle with
the electrode in it). Keep a record of the temperature calibrations in a lab notebook.
19.3 | Pre-Analysis Calibration
[Note: The pH buffer solutions are not used for any quantitative purpose. They are used to standardize the
electrode and as a diagnostic to verify that the pH measurement system is working as expected before
beginning analysis of the samples.]
19.3.1 | Use a pH Analytical Log Form to record all data (see Figure 17). |  | While still in standby mode,
reconnect the electrode lead at the back of the pH meter.
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-29
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
19.3.2 | Fill three 4 mL cups with pH 7 buffer. |  | Withdraw the electrode from the 4 M KCl bottle and wipe
the tip gently with a Kimwipe® to remove the bulk of the solution. |  |  | Rinse the electrode with one cup of pH 7
buffer. | Do not test pH of the first cup.
19.3.3 | Immerse the electrode in the second cup of the pH 7 buffer. |  |  | Use a small bottle or other support to
hold the cup up to the electrode while waiting for the meter reading to equilibrate.
19.3.4 Test the pH by turning to the pH mode of the meter. |  |  | Allow the reading to stabilize for at least 30 s.
Record the result on the log for "1st cup."
19.3.5 | Turn to standby mode, and then test the last cup of pH 7 buffer. |  |  | Record the results on the log for
the "2nd cup." | If the pH value for the 2nd cup is not 7.00 ± 0.01, adjust the "calib." knob to obtain a reading
of 7.00. | Note this adjustment on the log.
19.3.6 | Fill three 4 mL cups with pH 4 buffer. |  | With the meter in the standby mode, remove the cup
containing pH 7 buffer, wipe the tip of the electrode gently with a Kimwipe® and then rinse the electrode with
the first cup of pH 4 buffer.
19.3.7 | Test the next two cups of pH 4 buffer as above, recording the results on the log. |  |  | If the pH value
for the 2nd cup is not 4.00 ± 0.01, adjust the "slope" knob to get a reading of 4.00. |  |  |  | If the value for the second
cup was not 4.00 ± 0.03, the calibrations at pH 7 and at pH 4 must both be repeated.
19.4 | pH Test 0.01 N HClO4 | Solution
[Note: | The 0.01 N HClO 4 solution is used to prepare the ES solution, which is used to prepare the EA
solution. The pH value for the EA solution must be 4.09 ± 0.04. If this pH value is not achieved, the 0.01
N HClO 4 solution must be reprepared.]
19.4.1 | Calibrate the pH meter with pH 4 buffer.
19.4.2 | Rinse the pH electrode with DDW. |  | Wipe the tip of the electrode with a Kimwipe®.
19.4.3 | Fill three 4 mL cups with EA solution. |  | Measure the pH of the test EA solution as with the buffer
solutions this value must be 4.09 ± 0.10.
19.4.4 | If the above pH value is not achieved, follow the steps outlined in Sections 18.1.3 through 18.1.6
to reprepare the solutions. | Test the pH of the new solutions. |  | Repeat as necessary to obtain a pH of 4.09 ±
0.04.
19.4.5 | Leave the electrode immersed in the "2nd cup" with the meter in the standby mode until ready to
start analysis of the working standards.
19.5 | Analysis of Working Standard
[Note: Immediately following the EA analysis, start testing the working standards.]
19.5.1 With the pH meter still in the standby mode, remove the last cup from the electrode, gently wipe the
tip with a Kimwipe®, and immerse the electrode into the working standard cup #1.
[Note: | Only two cups are available for each working standard (also for filter extracts). |  |  | Thus, pH
measurement is made for both of the two cups for each sample. Also, the electrode tip is not wiped between
the 1st and 2nd cups of each sample.]
19.5.2 | After testing the pH of cup #1, test cup #1-. |  | Record the results of both on the pH Analytical Log
Form.
19.5.3 | With the meter in the stand-by mode, remove the #1- sample cup, wipe the electrode with a
Kimwipe®, and test one 2 mL cup of EA solution. |  |  | Rinse with DDW.
Page 4.2-30 | Compendium of Methods for Inorganic Air Pollutants |  |  | June 1999
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
19.5.4 Test a 2nd cup of EA solution; record the results for both cups on the logsheet. |  |  | Discard the 1st cup
of EA, but retain the 2nd cup to be used as the 1st cup for the next EA test.
19.5.5 Continue testing the remainder of the working standards, #1*, 1*-, ... 7, 7-, 7*, 7*-. Remember that
the electrode tip is wiped before and after each pair of test solutions, but not in between two cups of the same
sample.
[Note: | If there is trouble obtaining constant pH values, use a magnetic stirrer to keep the contents to be
measured uniform. If employed, ensure that the sample cups are insulated from any temperature increase of
the stirring platform, which may occur during extended use.]
19.5.6 |  | Use the mode control knob in the "temp." position to measure the temperature of the test solutions
every 5-10 samples and record the results on the logsheet.
19.6 | Analysis of Filter Extracts
After measuring the pH of the working standards, measure the pH of the filter extracts to which 20 FL of 2M
KCl solution (19.1.2) has been added, and record all results on the log. After all the filter extracts have been
tested make an additional test with the EA solution. At the end make a final test of pH 4 buffer. With the
mode control in the standby mode, shut down the pH meter by disconnecting the electrode lead at the back
of the meter, leaving the meter power cord plugged into the AC line. Immerse the electrode tip in the bottle
of 4 M KCl.
20. Atmospheric Species Concentration Calculations
The system described in the previous sections collects nitric acid (HNO 3), sulfur dioxide (SO 2), ammonia
(NH3), particulate sulfate (SO 4= ), particulate nitrate (NO 3- ), particulate ammonium (NH4+ ), and acidic (H + )
particles. The collection of each of these species is illustrated in Figure 1. |  |  | Nitric acid and sulfur dioxide
gases are collected on denuder one. Ammonia gas is collected on denuder two. Particulate sulfate, nitrate,
and ammonium are collected on the first (Teflon)® filter, while some of the particulate nitrate collected on
the Teflon® filter evaporates and collects on the second (nylon) filter. Also collected on the Teflon® filter
are fine particles that contain hydrogen ions (H+ ), though probably not free H + . Hydrogen ions are most
likely present in the H 3O + form. The concentration of these H+ ions indicates the atmospheres acid aerosol
content and is collected on the nylon filter. |  | Prepare the Teflon® filter extracts for pH analysis prior to IC
analysis for the particulate sulfate contents. Special precautions must be taken to prevent contamination of
the Teflon® filters by ammonia before either of the analyses.
20.1 | Assumptions of the Annular Denuder System
20.1.1 |  | A number of assumptions are made about the performance of the annular denuder system so that
validity of the calculations presented later in this section will hold true. |  | As discussed in Section 6, significant
interferences need to be considered to make accurate estimations of species concentrations. |  |  | The assumptions
are as follows:
| • | The first denuder stage collects 100% of sampled HNO 3 as nitrate. (Since the diffusivity of HNO3
is high, diffusion to the side walls is assumed to be very quick.)
| • | The first denuder collects 100% of the SO 2 as sulfite, which can oxidize to sulfate.
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-31
Method IO-4.2 | [5 empty] | Chapter IO-4
Acidic/Basic Constituents | [5 empty] | Atmospheric Acidic
[Note: Before analysis, add hydrogen peroxide (H 2O 2) to oxidize the sulfite (SO 3- ) to sulfate (SO 4= ) to
simplify the calculations.]
| • | The second denuder stage collects 100% of the sampled ammonia (NH 3) as ammonium ion (NH 4+ ).
| • | The Teflon® filter is 100% efficient for particulate sulfate, nitrate, and ammonia. Particle losses
are less than 1% on each denuder. This assumption may or may not stand true depending on the
|  | concentrations of the components in the air sampled. |  |  |  | Modifications may be needed to avoid low
|  | (or underestimates of) acidic measurements. |  |  | For example, another filter stage may need to be
added to accurately account for the particulate ammonia content of the air sampled. If ammonium
nitrate (NH4NO3) was collected on the Teflon® filter, its probability of evaporation is high.
Therefore, a citric acid-impregnated filter downstream would correct for the loss from the Teflon®
filter. Also, interaction of ammonia and sulfuric acid neutralizes the filter and causes the acidic
measurement to be biased. (Again diffusion rules the particle loss assumption; particles have lower
diffusivities than gases.)
| • | The nylon filter collects any nitrate that evaporates from the Teflon® filter.
20.1.2 |  | Analytical results from the IC and technician autoanalysis are given in units of μg/mL from NO3- ,
SO4=, and NH 4+, respectively.
20.1.3 |  | The extraction volume was 11 mL for Na2 CO3 |  |  | denuder(s) and 10 mL (i.e., 0.010 L) for the citric
acid denuder extractions and the nylon filter and 6.2 mL for the Teflon® filter.
20.2 | Calculation of Air Volume Sampled, Corrected to Standard Conditions
20.2.1 The total sample air volume, Vt , for each sample is calculated using the data from the Field Test Data
Sheet. | These data include the initial and final elapsed times, the initial rotameter reading, and the rotameter I.D.
No. | Use the calibration curve for the given rotameter to calculate the flow for the sample, in LPM, if
applicable. |  | Calculate the value of Vt |  | as follows:
V t = [F] [t]
where:
F = flow from the calibration curve, L per minute.
t = net elapsed time, min.
V t = total sample volume, L.
20.2.2 |  | Convert L to m3 | by:
V s = V t x (10 -3)
where:
V s = total sampling volume, m 3.
10 -3 = conversion factor, m 3/L.
20.2.3 |  | Calculate the air volume sampled, corrected to EPA-reference conditions:
Page 4.2-32 |  | Compendium of Methods for Inorganic Air Pollutants |  |  |  | June 1999
Chapter IO-4 | [10 empty] | Method IO-4.2
Atmospheric Acidic | [10 empty] | Acidic/Basic Constituents
[7 empty] | T | P
|  |  |  | V |  | V | Y(std | )(bar | )
|  |  |  | std | ' | s
[7 empty] | T | P
[7 empty] | m | std
where:
| V |  | volume of sample at EPA-reference conditions, m3 .
std
| Vs = |  | volume of gas sample through the dry gas meter, or calculated volume sampled as
indicated by rotameter (see Section 20.2.2), m3 .
| T std = |  | absolute EPA-reference temperature, 298EK.
| T m = |  | average flowmeter or dry gas meter temperature, EK.
| P bar |  | barometric pressure of flow or volume measurement condition, mm Hg.
| P std = |  | EPA-reference barometric pressure, 760 mm Hg.
| Y = |  | dry gas meter calibration factor (if applicable), dimensionless.
20.3 | Calculations of Concentration Using Results from IC and Technicon Autoanalysis
20.3.1 | Analytical Results for HNO 3 |  |  | from Na 2 CO 3 |  |  | Denuder
Cg (HNO 3 ), μg/m3
This factor 1.016 represents the ratio of molecular weights of HNO3- and NO 3- . Subscript "g" denotes "gas."
20.3.2 | Analytical Results for SO 2 |  |  | from Na 2 CO 3 |  |  | Denuder
Cg (SO 2 ), μg/m3
The factor 0.667 represents the ratio of molecular weights of SO2 | [9 empty] | and SO 4=. | Subscript "g" denotes "gas."
20.3.3 | Analytical Results for NH 3 |  |  | from Citric Acid Denuder
Cg (NH 3 ), μg/m3
This factor 0.944 represents the ratio of molecular weights of ammonia to ammonium ion. | [10 empty] | Subscript "g"
denotes "gas."
20.3.4 | Analytical Results for Particulate SO4= Captured on Teflon® Filter
|  |  | Cp | (SO 4 ), μg/m3 |  | [SO4= (μg/mL) x 6.2 mL]/Vstd
This formula expresses the assumption that essentially all of the particulate sulfate is collected on the
Teflon® filter and no evaporation occurs. |  |  |  | The subscript "p" denotes "particles."
20.3.5 | Analytical Results for Particulate NH4+ Captured on Teflon® Filter
|  |  | Cp | (NH 4+), μg/m3 |  | [NH4+ (μg/mL) x 6.2 mL]/Vstd
June 1999 |  |  | Compendium of Methods for Inorganic Air Pollutants | [7 empty] | Page 4.2-33
Method IO-4.2 | [11 empty] | Chapter IO-4
Acidic/Basic Constituents | [10 empty] | Atmospheric Acidic
This formula expresses the assumption that essentially all of the particulate ammonia is collected on this
Teflon® filter and no evaporation occurs. |  |  |  |  | The subscript "p" denotes "particles."
20.3.6 | Analytical Results for Particulate NO3G Captured on Teflon® Filter
|  | Cpt | (NO 3G p ), μg/m3 |  |  |  | [NO3 G (μg/mL) x 6.2 mL]/Vstd
Nitrates evaporator from the Teflon® filter during sampling, so one must also calculate the NO3G on the
nylon filter. | The subscript "pt" denotes particles on Teflon® filters.
|  | Cpn | (NO 3 G), μg/m3 |  |  |  | [NO3 G (μg/mL) x 10 mL]/Vstd
The subscript "pn" denotes particles on the nylon filter. | [7 empty] | Calculate the particulate nitrate concentration by:
|  |  | Cp | (NO 3 G) = Cpt |  |  | (NO 3 G) + Cpn |  | (NO 3 G)
20.4 | Calculations Using Results from pH Analysis
Earlier pH determinations have been based on the pH buffer concentrations, the activity of the solution, and
the antilog of the measured pH value. |  |  |  | More recent studies have steered away from the issue of activity by
comparing the results of the standards, thus alleviating errors introduced by basing the activities of ions retained
on filters on those retained in solution. The methodology developed from these more recent studies is described
herein. | The end results are reported in terms of mass of equivalent of ions. | [8 empty] | Appropriate values of accuracy
and precision with respect to H+ concentration for this method are 10% and 5%, respectively, for sample pH
values in the 4.00 to 7.00 range.
20.4.1 | Adjustment | for | Filter | vs. | Non-Filter |  | Standards. |  | This | adjustment | is | necessary | because
experiments showed that the measured acid concentration from filters doped with H 2 SO4 stock standards yielded
concentrations, as measured by the difference from EA solution, that were about 3% lower than the values
found for working standards (prepared without filters from the same stock standards). | [10 empty] | The results gave the
following relation (by linear regression):
Cf
where:
Cf
Cnf | [7 empty] | based on standards prepared without filters.
For each working standard (non-filter), on a given analysis day, calculate the "apparent net strong acid
concentration of H2 SO4 " as follows:
[7 empty] | -pHWS | -pHEA
[5 empty] | Cnf |  | - 10
Page 4.2-34 |  | Compendium of Methods for Inorganic Air Pollutants | [9 empty] | June 1999
Chapter IO-4 | [11 empty] | Method IO-4.2
Atmospheric Acidic | [11 empty] | Acidic/Basic Constituents
where:
pHWS = measured pH for a working standard (or apparent strong acid concentration for non-filter H 2 SO4
doped standards).
pHEA = measured pH for the EA solution (or apparent strong acid concentration for non-filter, non-H2 SO4
doped standard).
After calculating the Cnf |  |  | values for each working standard, use the above equation to calculate the adjusted
values of Cf |  | for each.
| 20.4.2 | Determination of Nominal Strong Acid Concentration for Filter Samples. | [9 empty] | The apparent net
strong acid concentration of each sample filter extract, Cs, is calculated as with the working standards:
[8 empty] | -pHS | -pHEA
[7 empty] | Cs = 10 | - 10
where:
pHS = measured pH of the sample filter extract (or apparent strong acid concentration for sample filters
extracts).
| pHEA = measured |  | pH | for | the | EA | solution | (or | apparent | strong | acid | concentration | for | non-filter,
|  | non-H 2 SO4 | standards).
[Note: The Cs values for the filter extracts are directly comparable to the C f values for the working standards,
since the Cf values have been adjusted for the difference in apparent acid concentration for tests made with
filters and tests made without filters.]
| 20.4.3 Calculation of Strong Acidity Aerosol Concentration. | [8 empty] | Calculate the final concentration of
apparent net fine particle (
[7 empty] | C(H +) = Cf |  | / Vstd
where:
CH + = apparent net fine particle str\\\\\=ong acidity concentration, Fg/m3 .
Cf
Vstd = volume of sampled gas at EPA-referenced conditions (see Section 17.2.3), m3.
21. | Variations of Annular Denuder System Usage
As described in Sections 3 and 4, the ADS is used to measure reactive acidic (SO2 and HNO 3 ) and basic (NH3 )
gases, particulate sulfates and nitrates, and strong acidity of atmospheric particles (
air. | The unique features of the ADS that separate it from established air monitoring methods are the ability of
sampling artifacts to be eliminated from the collected gases and particles and preservation of the samples for
subsequent analysis, which is accomplished by removing NH3 | [9 empty] | in the gas stream with a citric acid coated
denuder, thus reducing the probability of the particulate acid sulfates (SO4=) captured on the Teflon® filter from
being neutralized. The ADS configuration described in this methodology clearly illustrates these unique features.
The elutriator is designed to allow only particles with | [11 empty] | The impactor is
designed to reduce the possibility of coarse particle infiltration even further. | [10 empty] | And finally, the sequence of the
June 1999 |  |  | Compendium of Methods for Inorganic Air Pollutants | [9 empty] | Page 4.2-35
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
denuders reduces interference of possible chemical reactions that could cause under-or over-estimations of
concentrations to be made. |  | Although this configuration is recommended for measuring these gases and
particulates, the user may wish to measure only one or two of the chemical species. The following discussion
will present possible variations of the ADS to accommodate such usages.
21.1 Today, the ADS is being used in intercomparison studies to assess NH3 concentration differences indoors
and outdoors. | The assembly used here consists of an elutriator-impactor assembly, an annular denuder, and
a filter pack assembly. |  | The elutriator-impactor assembly and the annular denuder are both smaller than those
described earlier. The filter pack is available in the smaller size, but an adaptor is also available to assemble the
smaller annular denuder to the larger filter pack assembly. |  | This system is referred to as the personal sampler
(see Figure 18). | It is designed for sampling while attached to the shirt of a worker. |  | The personal sampler can
be used to measure other chemical species in indoor air by simply changing the reactive surface (coating) of the
annular denuder or the types of filters used.
21.2 Another variation of ADS application is simultaneous use in parallel with a fine particle sampler. The fine
particle sampler assembly is very similar to the annular denuder assembly. |  | The main difference is that a
flow-straightener tube replaces the annular denuder. The flow-straightener is a shorter version, 1-1/4 to 4" long,
of the annular denuder and creates even air flow across the filters to collect particulate matter. |  |  | Again the
elutriator-impactor assembly and flow-straightener are available in smaller sizes with accommodating filter pack
assemblies. | In addition, the ADS carrying and shipping cases as well as the sampling box can be adjusted to
accommodate the ADS and fine particle sampler. |  | The assemblies as they would appear in the sampling box
ready for sampling are illustrated in Figure 19.
21.3 If one has interest in quantitations HNO2 utilizing the ADS, special sampling and analytical concerns must
be addressed. | As identified in Section 24, Citation 14, guidance is given for accurate quantitation of HNO2 |  | in
ambient air.
21.4 | The annular denuder methodology has been extended to other constituents, as indicated in Table 5.
22. | Method Safety
This procedure may involve hazardous materials, operations, and equipment. |  | This method does not purport
to address all of the safety problems associated with its use. |  | The user must establish appropriate safety and
health practices and determine the applicability of regulatory limitations prior to the implementation of this
procedure. | These practices should be part of the user's SOP manual.
23. | Performance Criteria and Quality Assurance (QA)
Required quality assurance measures and guidance concerning performance criteria that should be achieved
within each laboratory are summarized and provided in the following section.
Page 4.2-36 |  | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
23.1 | Standard Operating Procedures (SOPs)
23.1.1 | SOPs should be generated by the users to describe and document the following activities in their
laboratory: (1) assembly, calibration, leak check, and operation of the specific sampling system and equipment
used; (2) preparation, storage, shipment, and handling of the sampler system; (3) purchase, certification, and
transport of standard reference materials; and (4) all aspects of data recording and processing, including lists of
computer hardware and software used.
23.1.2 | Specific instructions should be provided in the SOPs and should be readily available to and
understood by the personnel conducting the monitoring work.
23.2 | QA Program
The user should develop, implement, and maintain a quality assurance program to ensure that the sampling
system is operating properly and collecting accurate data. Established calibration, operation, and maintenance
procedures should be conducted regularly and should be part of the QA program. |  |  | Calibration procedures
provided in Sections 17 and 19, operation procedures in Sections 14 and 17, and maintenance procedures in
Section 17 of this method and the manufacturer's instruction manual should be followed and included in the QA
program. | Additional QA measures (e.g., trouble shooting) and further guidance in maintaining the sampling
system are provided by the manufacturer. |  | For detailed guidance in setting up a quality assurance program, the
user is referred to the Code of Federal Regulations (Section 24, Citation 12) and the U. S. EPA Handbook on
Quality Assurance (Section 24, Citation 13).
23.2.1 Field QA. |  | It is recommended that the flow rates of each denuder system be audited at least
quarterly.
23.2.2 Laboratory Quality Control (QC). |  | It is recommended that laboratory QC program include the
following as minimum requirements.
23.2.2.1 Include a reagent blank with each set of twenty (20) samples or less of each matrix (i.e., denuder
extract, filter extract, etc.).
23.2.2.2 Include a laboratory duplicate with each set of twenty (20) samples or less of each matrix.
23.2.2.3 Include a laboratory control sample (LCS), also known as a laboratory blank spike (LBS) with
each set of twenty (20) samples or less of each matrix.
24. References
1. | Waldman, J. M., Operations Manual for the Annular Denuder System Used in the U. S. EPA/RIVM
Atmospheric Acidity Study, UMPNJ - Robert Wood Johnson Medical School, Piscataway, NJ, August 28,
1987.
2. | American Chemical Society Subcommittee on Environmental Chemistry, "Guidelines for Data
Acquisition and Data Quality Evaluation in Environmental Chemistry," Analyt. Chem., Vol. 52:2242-2249,
1980.
3. | Sickles, II, J. E., Sampling and Analytical Methods Development for Dry Deposition Monitoring,
Research Triangle Institute Report No. RTI/2823/00-15F, Research Triangle Institute, Research Triangle
Park, NC, July 1987.
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-37
Method IO-4.2 | [9 empty] | Chapter IO-4
Acidic/Basic Constituents | [7 empty] | Atmospheric Acidic
4. Forrest, J., and Neuman, L., "Sampling and Analysis of Atmospheric Sulfur Compounds for Isotopic
Ratio Studies," Atmos. Environ., Vol. 7:562-573, 1973.
5. | Stevens, R. K., et al., ACGIH Symposium: |  |  | "Inlets, Denuders and Filter Packs to Measure Acidic
Inorganic Pollutants in the Atmosphere," Aislomer Conference Center, Pacific Grove, CA, February 16,
1986.
6. Appel, B. R., Povard, V., and Kothney, E. L., "Loss of nitric acid within inlet devices for Atmospheric
Sampling," Paper presented at 1987 U. S. EPA/APCA Symposium: Measurement of Toxic and Related Air
Pollutants, Research Triangle Park, NC, 3-6 May 1987.
7. Braman, R. S., et al., "Tungstic Acid for Preconcentration and Determination of Gaseous and Particulate
Ammonia and Nitric Acid in Ambient Air," Analyt. Chem., Vol. 54:358-364, 1983.
8. | Ferm, M., Concentration Measurements and Equilibrium Studies of Ammonium, Nitrate and Sulphur
Species in Air and Precipitation, Doctoral Thesis, Department of Inorganic Chemistry, Goteborg University,
Goteborg, Sweden, 1986.
9. Ferm, M., and Sjodin A., "A Sodium Carbonate Coated Denuder for Determination of Nitrous Acid in
the Atmosphere," Atmos. Environ., Vol. 19:979-985, 1985.
10. Riggin, R. M., Technical Assistance Document for Sampling and Analysis of Toxic Organic Compounds
in Ambient Air, U.S. Environmental Protection Agency, EPA Office of Research and Development,
Atmospheric Research and Exposure Assessment Laboratory, Research Triangle Park, NC, 27511,
EPA-600/4-83-027, November, 1983.
11. | Stevens, R. K., and Rickman, E., Jr., "Research Protocol/Method for Ambient Air Sampling with
Annular Denuder Systems," U.S. Environmental Protection Agency, EPA Office of Research and
Development, Atmospheric Research and Exposure Assessment Laboratory, Research Triangle Park, NC,
27511, ASRL-ACPD-RPM 003, January 1988.
12. 40 CFR Part 58, Appendix A, B.
13. Quality Assurance Handbook for Air Pollution Measurement Systems, Volume II - Ambient Air Specific
Methods | (Interim | Edition), | U.S. | Environmental | Protection | Agency, | EPA | Office | of | Research | and
Development, Atmospheric Research and Exposure Assessment Laboratory, Research Triangle Park, NC,
27511, EPA-600/R-94/038b, May, 1972.
14. Febo H., Perino, C., and Cortiello, M., "A Denuder Technique for the Measurement of Nitrous Acid
in Urban Atmospheres," Atmos. Environ., Vol. 27:1721-1728, 1993.
15. Winberry, W.T., Jr., "Determination of the Strong Acidity of Atmospheric Fine Particles (
Using Annular Denuder Technology," U.S. Environmental Protection Agency, EPA Office of Research and
Development, Atmospheric Research and Exposure Assessment Laboratory, Research Triangle Park, NC,
27511, EPA-600/R-93-D37, November 1992.
Page 4.2-38 |  | Compendium of Methods for Inorganic Air Pollutants | [7 empty] | June 1999
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
TABLE 1. ESTIMATED DETECTION AND QUANTIFICATION LIMITS
FOR THE ANNULAR DENUDER SYSTEM1
Sampling period
Detection limits, μg/m3 | 1 hour | 1 day | 1 week
Gaseous species
SO2 | 3.1 | 0.13 | 0.02
HNO3 | 2.0 | 0.08 | 0.01
NH 3 | 5.6 | 0.25 | 0.04
Particulate matter
SO4= | 1.6 | 0.07 | 0.01
NO3- | 1.8 | 0.08 | 0.01
H + | 1.0 | 0.05 | 0.001
Quantification limits, Fg/m3
Gaseous species
SO2 | 10.4 | 0.43 | 0.06
HNO3 | 6.8 | 0.28 | 0.04
NH 3 | 20.0 | 0.83 | 0.12
Particulate matter
SO4 = | 5.3 | 0.22 | 0.03
NO3- | 6.1 | 0.25 | 0.04
H + | 2.1 | 0.15 | 0.01
1Samples analyzed by ion chromatography. | Detection limits are taken as three standards deviations above field
blanks. | Quantification limits are taken as ten standard deviations above field blanks. |  | Both the detection | and
quantification limits were estimated assuming that the variance is independent of concentration.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-39
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
TABLE 2. ACCELERATOR JET DIAMETERS AND CORRESPONDING REYNOLDS
NUMBER (RE) FOR SELECTED FLOW RATES TO OBTAIN 2.5 μm DIAMETER
(AERODYNAMIC) PARTICLES SEPARATION
Flow rate,
| L/min | Jet diameter, mm | Reynolds number (RE)
| 1.0 | 1.55 | 900
| 2.0 | 1.97 | 1,400
| 5.0 | 2.65 | 2,700
| 10.0 | 3.33 | 4,200
| 12.0 | 3.55 | 4,700
| 15.0 | 3.85 | 5,500
| 16.7 | 4.00 | 6,000
| 20.0 | 4.25 | 6,600
Page 4.2-40 | Compendium of Methods for Inorganic Air Pollutants |  | June 1999
Chapter IO-4 |  |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
TABLE 3. SUMMARY OF KEY PROBE SITING CRITERIA FOR ACID AEROSOL
MONITORING STATIONS
Factor |  |  | Criteria
Vertical spacing above ground | • | Representative of the breathing zone and avoiding effects of
|  | obstruction, obstacles, and roadway traffic. |  | Height of probe
intake above ground in general, 2-3 m above ground and 2-15 m
above ground in the case of nearby roadways.
| • | About 1 m or more above the structure where the sampler is
located.
Horizontal spacing from | • | Minimum horizontal separation from obstructions such as trees
obstruction and obstacles |  | should be > 20 m from the dripline and must be 10 m from the
dripline when the trees act as an obstruction.
| • | Distance from sampler inlet to an obstacle such as a building must
be at least twice the height the obstacle protrudes above the
sampler.
| • | If a sampler is located on a roof or other structures, there must be
a minimum of 2 m separation from walls, parapets, penthouses,
etc.
| • | There must be sufficient separation between the sampler and a
|  | furnace or incinerator flue. | The separation distance depends on
the height and the nature of the emissions involved.
Unrestricted airflow | • | Unrestricted airflow must exist in an arc of at least 270 degrees
around the sampler, and the predominant wind direction for the
monitoring period must be included in the 270 degree arc.
| • | A sufficient separation must exist between the sampler and nearby
Spacing from roads |  | roadways to avoid the effect of dust re-entrainment and vehicular
emissions on the measured air concentrations.
| • | Sampler should be placed at a distance of 5-25 m from the edge of
the nearest traffic lane on the roadway depending on the vertical
placement of the sampler inlet which could be 2-15 m above
ground.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  |  | Page 4.2-41
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
TABLE 4. DILUTION RATES
|  | Volume of | Equivalent
|  | 1.0 N H2SO 4 | Strong Acid
| Standard H 2SO 4 | Added to Each | Mass Collected
| Flask No. | Flask, mL | on Filter, Fg a | Approximate pH
| 1 | 0.000 | 0 | 4.09
| 2 | 0.025 | 4.90 | 4.01
| 3 | 0.050 | 9.80 | 3.95
| 4 | 0.100 | 19.60 | 3.84
| 5 | 0.200 | 39.20 | 3.68
| 6 | 0.400 | 78.40 | 3.48
| 7 | 0.800 | 156.80 | 3.23
aBased on 6.2 mL extraction volume.
TABLE 5. OTHER CHEMICAL SPECIES
AMENDABLE TO ANNULAR DENUDER TECHNOLOGY
Chemical Species
Collected on Annular
Symbol | Denuder |  | Coating | Comments
O 3 | Ozone |  | 1% (wt/wt) | Analysis by IC or
|  | NaHCO3/glycerin |  | colorimetry
HCl | Hydrogen Chloride | Na2CO3-alkali coating |  | Extracted with water/
HF | Hydrogen Fluoride |  |  | analyzed by IC
PAHs | Semi-volatile polycyclic |  | Ground | Soxhlet
SVOCs | aromatic hydrocarbons | XAD-4/glycerin |  | extraction/GC/MS
analysis
Page 4.2-42 | Compendium of Methods for Inorganic Air Pollutants |  |  | June 1999
Chapter IO-4 | Method IO-4.2
Atmospheric Acidic | Acidic/Basic Constituents
es
Figure 1. Schematic view of Annular Denuder showing species collected.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-43
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
Figure 2. Annular Denuder system in field sampling box.
Page 4.2-44 | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic | Acidic/Basic Constituents
Figure 3. Available elutriator and acceleration jet assemblies.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-45
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents | Atmospheric Acidic
Figure 4. Glass Annular Denuder with inset impactor assembly.
Page 4.2-46 | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
Figure 5. Side view impactor/coupler assembly.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-47
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
Figure 6. Schematic view of Annular Denuder with cyclone
adaptor for removal of coarse particles.
Page 4.2-48 | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
Figure 7. Internal schematic of Annular Denuder.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-49
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
Figure 8. Filter pack assembly.
Page 4.2-50 | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
Figure 9. Drying train and manifold.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-51
Method IO-4.2 | Chapter IO-4
Acidic/Basic Constituents | Atmospheric Acidic
Figure 10. Annular Denuder system with cyclone in heated sampling case.
Page 4.2-52 | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
Figure 11. Frit removal from pin.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-53
Method IO-4.2 | Chapter IO-4
Acidic/Basic Constituents | Atmospheric Acidic
Figure 12. Side view impactor/coupler assembly with disc removal tools.
Page 4.2-54 | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  |  |  |  | Method IO-4.2
Atmospheric Acidic |  |  |  | Acidic/Basic Constituents
GENERAL
Project: |  |  | Date:
Site: |  |  | Location of Sampler:
Location:
Operator:
Sample Code:
| EQUIPMENT |  |  | Sampler
Mass Flow
Controller No.: |  |  | Sodium Carbonate Denuder No.
Lab Calibration Date: |  |  | Citric Acid Denuder No.:
Flow Rate Set Point: |  |  | Filter Assembly No.:
Calibrated By:
Rotameter No.:
DGM No.:
| SAMPLING DATA |  |  | Stop
Time
Time:
Flow Rate:
Temperature:
Pressure:
Avg. Flow Rate:
Leak Check (Before):
(After):
Total Sample Vol.:
Flow Maintained Rate: |  | (±5%)
| Flow | Ambient | Barometric | Relative
| Rate (Q), | Temperature, | Pressure, | Humidity, %
Time | L/min | EC | mm Hg |  | Comments
Figure 13. Annular Denuder Field Test Data Sheet.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  |  |  | Page 4.2-55
Method IO-4.2 | Chapter IO-4
Acidic/Basic Constituents | Atmospheric Acidic
Analytical
Figure 14. Major components of a commercially available ion chromatographer.
Page 4.2-56 | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic | Acidic/Basic Constituents
Figure 15. Chromatograms of denuder/filter extract performed by the ion chromatography.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-57
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents | Atmospheric Acidic
Figure 15 (Cont'd). Chromatograms of denuder/filter extract performed by the ion chromatography.
Page 4.2-58 | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 | Method IO-4.2
Atmospheric Acidic | Acidic/Basic Constituents
Figure 16. Technicon autoanalyzer flow diagram for ammonia analysis.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-59
Method IO-4.2 | [7 empty] | Chapter IO-4
Acidic/Basic Constituents | [7 empty] | Atmospheric Acidic
Determination of the Strong Acidity of Atmospheric Fine-Particles (
| Name: |  |  |  | Sample I.D.:
| Date: |  |  |  | Location:
LAB:
RUN NUMBER
| Constituent | 1 | 2 | 3 | 4 | 5 | 6 | 7
pH 7 Buffer
1
2
3
pH 4 Buffer
1
2
3
EA Solution
1
2
3
Working Standards
1A1
1A2
EA
1B1
1B2
EA
2A1
2A2
Temp.
EA
2B1
2B2
EA
3A1
3A2
EA
Temp.
3B1
3B2
EA
Figure 17. pH Analytical Laboratory Log Form
Page 4.2-60 |  | Compendium of Methods for Inorganic Air Pollutants | [5 empty] | June 1999
Chapter IO-4 | [6 empty] | Method IO-4.2
Atmospheric Acidic | [5 empty] | Acidic/Basic Constituents
RUN NUMBER
Constituent | 1 | 2 | 3 | 4 | 5 | 6 | 7
4A1
4A2
EA
4B1
4B2
Temp.
EA
5A1
5A2
EA
5B1
5B2
EA
6A1
6A2
Temp.
EA
6B1
6B2
EA
EA
7A1
7A2
EA
7B1
7B2
Temp.
Sample Extracts
A
A1
B
B1
C
C1
D
D1
E
E1
EA
Temp.
EA Solution
1
2
3
pH4 Buffer
1
2
3
Figure 17 (cont). pH Analytical Laboratory Log Form
June 1999 | Compendium of Methods for Inorganic Air Pollutants | [5 empty] | Page 4.2-61
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
Figure 18. Field Particle personal sampler.
Page 4.2-62 | Compendium of Methods for Inorganic Air Pollutants | June 1999
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic | Acidic/Basic Constituents
Figure 19. Annular Denuder system with parallel fine particle sampler.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-63
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
This page intentionally left blank
Page 4.2-64 | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 1
Name: “Table 1 of 1 on page 1 of 68”
Table: 1
EPA/625/R-96/010a
Compendium of Methods
for the Determination of
Inorganic Compounds
in Ambient Air
Compendium Method IO-4.2
DETERMINATION OF REACTIVE
ACIDIC AND BASIC GASES
AND STRONG ACIDITY OF
ATMOSPHERIC FINE PARTICLES
(
Center for Environmental Research Information
Office of Research and Development
U.S. Environmental Protection Agency
Cincinnati, OH 45268
June 1999

## Page 2
Name: “Table 1 of 1 on page 2 of 68”
Table: 2
Method IO-4.2
Acknowledgments
This Method is a part of Compendium of Methods for the Determination of Inorganic Compounds in Ambient
Air (EPA/625/R-96/010a), which was prepared under Contract No. 68-C3-0315, WA No. 2-10, by Midwest
Research Institute (MRI), as a subcontractor to Eastern Research Group, Inc. (ERG), and under the
sponsorship of the U.S. Environmental Protection Agency (EPA). |  | Justice A. Manning, John O. Burckle,
Scott R. Hedges, Center for Environmental Research Information (CERI), and Frank F. McElroy, National
Exposure Research Laboratory (NERL), all in the EPA’s Office of Research and Development, were
responsible for overseeing the preparation of this method. |  | Other support was provided by the following
members of the Compendia Workgroup:
• | James L. Cheney, U.S. Army Corps of Engineers, Omaha, NE
• | Michael F. Davis, U.S. EPA, Region 7, KC, KS
• | Joseph B. Elkins Jr., U.S. EPA, OAQPS, RTP, NC
• | Robert D. Lewis, U.S. EPA, NERL, RTP, NC
• | Justice A. Manning, U.S. EPA, ORD, Cincinnati, OH
• | William A. McClenny, U.S. EPA, NERL, RTP, NC
• | Frank F. McElroy, U.S. EPA, NERL, RTP, NC
• | William T. "Jerry" Winberry, Jr., EnviroTech Solutions, Cary, NC
This Method is the result of the efforts of many individuals. Gratitude goes to each person involved in the
preparation and review of this methodology.
Author(s)
• | William T. "Jerry" Winberry, Jr., EnviroTech Solutions, Cary, NC
• | Thomas Ellestad, U.S. EPA, RTP, NC
• | Bob Stevens, U.S. EPA, RTP, NC
Peer Reviewers
• | Delbert Eatough, Brigham Young University, Provo, UT
• | Shere Stone, University Research Glassware Corp., Chapel Hill, NC
• | Petros Koutrakis, Harvard School of Public Health, Boston, MA
• | J. Waldman, Robert Wood Johnson Medical School, New Brunswick, NJ
• | Lauren M. Drees, U.S. EPA, NRMRL, Cincinnati, OH
DISCLAIMER
This Compendium has been subjected to the Agency's peer and administrative review, and it has
been approved for publication as an EPA document. Mention of trade names or commercial products
does not constitute endorsement or recommendation for use.
ii

## Page 3
Name: “Table 1 of 1 on page 3 of 68”
Table: 3
Method IO-4.2
Determination of Reactive Acidic and Basic
Gases and Strong Acidity of Atmospheric
Fine Particles (
TABLE OF CONTENTS
Page
1. Scope . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  |  | 4.2-1
2. | Applicable Documents |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-2
| 2.1 | ASTM Standards . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-2
| 2.2 | Other Documents | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-2
3. | Summary of Method |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-2
4. | Significance | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-2
5. | Definitions . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-3
6. | Factors Affecting Denuder Efficiency |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-4
7. | Apparatus | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-4
| 7.1 | Sampling . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-5
| 7.2 | Analysis | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-7
8. | Reagents and Materials . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-8
9. | Preparation of Coating and Extraction Reagents |  |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . | 4.2-10
| 9.1 | Impactor Frit Coating Solution Preparation |  | . . . . . . . . . . . . . . . . . . . . . . . . . . | 4.2-10
| 9.2 | Impactor Frit Extraction Solution Preparation . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-10
| 9.3 | Annular Denuder Coating Solutions Preparation . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-10
10. | Elutriator and Acceleration Jet (Inlet) Assembly |  |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . | 4.2-10
11. | Impactor Frit Preparation and Installation |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-11
| 11.1 | Impactor Frit Installation | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-11
| 11.2 | Impactor Frit Preparation . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-11
12. | Filter Pack Preparation and Assembly . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-11
13. | Annular Denuder System Preparation |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-12
| 13.1 | Annular Denuder Coating Procedure . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-13
| 13.2 | Annular Denuder Drying Procedure . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-13
| 13.3 | Annular Denuder System (ADS) Assembly |  | . . . . . . . . . . . . . . . . . . . . . . . . . . | 4.2-13
| 13.4 | Laboratory Leak-Check of ADS . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-14
14. | Sampling . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-14
| 14.1 | Placement of Denuder System . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-14
| 14.2 | Start-Up | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-15
| 14.3 | Sample Shutdown | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-16
| 14.4 | Corrective Action for Leak Test Failure | . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-17
15. | ADS Disassembly |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-17
16. | Extraction Procedures |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-18
| 16.1 | Impactor Frit Coating Extraction (Optional) . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-18
| 16.2 | Denuder Extractions | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-18
| 16.3 | Filter Extraction . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-19
iii

## Page 4
Name: “Table 1 of 1 on page 4 of 68”
Table: 4
TABLE OF CONTENTS (continued)
Page
17. | Ion Chromatography Analysis |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-20
| 17.1 | Standards Preparation | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-21
| 17.2 | Reagent Preparation . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-21
| 17.3 | Sample Preparation . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-22
| 17.4 | Basic System Operations - Start-up and Shut-down |  | . . . . . . . . . . . . . . . . . . . . . | 4.2-23
| 17.5 | Basic Troubleshooting | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-25
18. | Ammonia Analysis by Technicon Autoanalysis |  | . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-26
| 18.1 | Standards and Stock Solutions Preparation . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-27
| 18.2 | Reagent Preparation . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-27
19. | pH Analysis | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-28
| 19.1 | Standard and Reagent Preparation | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-28
| 19.2 | Calibration of pH Meter . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-30
| 19.3 | Pre-Analysis Calibration . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-30
| 19.4 | pH Test 0.01 N HClO4 Solution . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-31
| 19.5 | Analysis of Working Standard . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-31
| 19.6 | Analysis of Filter Extracts | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-32
20. | Atmospheric Species Concentration Calculations . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-32
| 20.1 | Assumptions of the Annular Denuder System . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-32
| 20.2 | Calculation of Air Volume Sampled, Corrected to Standard Conditions . . . . . . . . |  |  | 4.2-33
| 20.3 | Calculations of Concentration Using Results from IC and Technicon
|  | Autoanalysis . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-34
| 20.4 | Calculations Using Results from pH Analysis . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-35
21. | Variations of Annular Denuder System Usage . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-36
22. | Method Safety . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-37
23. | Performance Criteria and Quality Assurance (QA) . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-37
| 23.1 | Standard Operating Procedures (SOPs) . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  | 4.2-38
| 23.2 | QA Program | . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  | 4.2-38
24. | References . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . |  |  |  | 4.2-38
iv

## Page 5
Name: “Table 1 of 1 on page 5 of 68”
Table: 5
Chapter IO-4
Atmospheric Basic & Acidic Constituents
Method IO-4.2
DETERMINATION OF REACTIVE ACIDIC AND BASIC
GASES AND STRONG ACIDITY OF
ATMOSPHERIC FINE PARTICLES (
1. Scope
1.1 | The quantitative measurement of reactive acidic and basic gases and strong acidity of atmospheric fine
particles in ambient air using annual denuder technology is described in this method. |  | The difference between
Inorganic Compendium Methods IO-4.1 and IO-4.2 is that the latter accounts for possible interference from
the dissociation of ammonium nitrate aerosol from particles collected on the filter by two mechanisms:
NH4NO3 = NH3(g) + HNO3(g)
NH4NO3 + NH4HSO 4 = (NH4) 2SO 4 + HNO3
Consequently, an accurate and quantitative value for determining strong acidity of atmospheric fine
particle is calculated.
1.2 The chemical species that can be determined by this method are gaseous SO2 , HNO 2 , HNO 3 , and NH 3 and
particulate SO=4, NO 3- , NH +4, and H +. | Detection and quantitation limits are given in Table 1.
1.3 | This method is a composite of methodologies developed by the U. S. Environmental Protection Agency
(EPA), Harvard University and the CNR Laboratories (Italy). A number of air pollution studies in Italy, United
States, Canada, Mexico, Germany, Austria, and Spain, and in public health services, epidemiology, and
environmental research centers have used this method.
1.4 The equipment described herein is used to measure acidic and basic gases and strong acidity of atmospheric
fine particles contained in ambient air. The methodology originally was developed for monitoring regional-scale
acidic and basic gases and strong acidity of atmospheric fine particle in support of EPA's field programs
involving the Integrated Air Cancer Research Program and the Acid Deposition Network. |  | Similarly, the
methodology has been used to characterize the urban haze in Denver, Houston, and Los Angeles.
1.5 | The techniques, procedures, equipment, and other specifications comprising this method are derived and
composited from those actually used by contributing research organizations and, therefore, are known to be
serviceable and effective. | At this stage, this method is a unified, consensus, tentative, draft method intended
for further application and testing. Users should be advised that the method has not yet been adequately tested,
optimized, or standardized. | Many of the specifications have been initially established by technical judgement
but have not been subjected to ruggedness testing. | In some cases, alternative techniques, equipment, or
specifications may be acceptable or superior. | In applying the method, users are encouraged to consider
alternatives, with the understanding that they should test any such alternatives to determine their adequacy and
confirm and document their advantages. Information and comments are solicited on improvements, alternative
equipment, techniques, specifications, performance, or any other aspect of the method.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-1

## Page 6
Name: “Table 1 of 1 on page 6 of 68”
Table: 6
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
2. Applicable Documents
2.1 | ASTM Standards
• D1356 Definitions of Terms Related to Atmospheric Sampling and Analysis.
2.2 | Other Documents
• Ambient Air Studies (1-14).
• U. S. EPA Acid Aerosol Document (15).
3. Summary of Method
3.1 | The annular denuder system (ADS) consists of (1) an inlet with an impactor or cyclone preseparator
designed to remove all particles with a diameter (aerodynamic) of 2.5 μm or greater, (2) annular denuders to
quantitate acidic and basic gases, and (3) a filter pack for atmospheric acidity and particles. |  |  | In operation,
ambient air is drawn through an elutriator-accelerator jet assembly, an impactor frit and coupler assembly, and
past glass denuder walls that have been etched and coated with chemicals that absorb the gaseous species of
interest. | The remaining air stream is then filtered through Teflon® and Nylasorb® membrane filters. Teflon®
and nylon membrane filters are used to capture ammonium and nitrate aerosol and sulfate particulate matter.
Nitric acid and sulfur dioxide will also be collected by the nylon filter but these measurements are treated as
interference. The ADS is illustrated in Figure 1. The field sampling box with the ADS and pump-timer system
is shown in Figure 2.
3.2 | Following each run, the ADS assembly is removed from its field housing, its ends are capped, and it is
brought back to the laboratory. In the laboratory, the assembly pieces are uncoupled and capped. The annular
denuders are extracted with 5 mL of deionized water. |  |  | The extracted solutions are subsequently analyzed for
ions corresponding to the collected gaseous species (see Figure 1). The filters are placed into filter bottles where
5 or 10 mL of the ion chromatographic (IC) eluent are pipetted into each filter bottle with the filters faced
downward and completely covered by |  | the eluent. | The filter bottle is capped and put in an ultrasonic bath for
30 min. | The bottles are stored in a clean refrigerator at 5EC until analysis.
3.3 | The analysis of anion and cation concentrations collected by the denuders and filter pack is typically
performed by ion chromatographic and Technicon® colorimeter autoanalytic procedures. The H + concentration
of extracts from the Teflon® filter downstream of the denuders is performed by pH measurements using
commercially available pH meters calibrated with standards.
4. Significance
4.1 | Reactive acidic (SO2 | and HNO 3 ) and basic (NH3 ) gases and particles are found in the atmosphere as a
result of emissions from a variety of fossil fuel combustion sources, including industrial and commercial
facilities, industrial processes, etc. Measurements of these chemical species are currently being used in a range
of environmental studies, such as in (1) epidemiological programs to assess the impact of acid aerosols on
respiratory impairment, (2) receptor modeling to determine the origin of particles that impact EPA's PM2.5 |  |  | or
Page 4.2-2 |  | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 7
Name: “Table 1 of 1 on page 7 of 68”
Table: 7
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
PM10 | air particulate standard, (3) assessment of the impact of particulate nitrate and sulfate on visibility, and
(4) quantification of the impact of acidic and basic air pollutants on issues related to acid rain.
4.2 Unique features of the annular denuder that separate it from other established monitoring methods are the
elimination of sampling artifacts due to interaction between the collected gases and particles and the preservation
of the samples for subsequent analysis. This preservation is accomplished by removing NH3 in the gas stream
by the citric acid coated denuder and reducing the probability of the particulate sulfate (SO4=) captured by the
filter pack from being neutralized to ammonium sulfate [(NH4 ) 2 SO4 ]. |  | If NH 3 | is not extracted from the gas
stream prior to filtration, particulate sulfate and gaseous sulfur dioxide correction would be required for accurate
measurements. |  | These biases in its configuration and analytical methodology are addressed in this method.
5. Definitions
Definitions used in this document and any user prepared Standard Operating Procedures (SOPs) should be
consistent with ASTM D1356. All abbreviations and symbols are defined within this document at the point
of use.
5.1 | Particulate Mass. | A generic classification in which no distinction is made on the basis of origin, physical
state, and range of particle size. (The term "particulate" is an adjective, but it is commonly used incorrectly as
a noun.)
5.2 | Primary Particles (or Primary Aerosols). | Dispersion aerosols formed from particles that are emitted
directly into the air and that do not change form in the atmosphere. |  | Examples include windblown dust and
ocean salt spray.
5.3 | Secondary Particles (or Secondary Aerosols). | Dispersion aerosols that form in the atmosphere as a
result of chemical reactions, often involving gases. A typical example is sulfate ions produced by photochemical
oxidation of SO2 .
5.4 | Particle. | Any object having definite physical boundaries in all directions, without any limit with respect
to size. | In practice, | the particle size range of interest is used to define "particle." | In atmospheric sciences,
"particle" usually means a solid or liquid subdivision of matter that has dimensions greater than molecular radii
(~10 nm); there is also not a firm upper limit, but in practice it rarely exceeds 1 mm.
5.5 | Aerosol. | A disperse system with a gas-phase medium and a solid or liquid disperse phase. | Often,
however, individual workers modify the definition of "aerosol" by arbitrarily requiring limits on individual
particle motion or surface-to-volume ratio. |  | Aerosols are formed by (1) the suspension of particles due to
grinding or atomization or (2) the condensation of supersaturated vapors.
5.6 | Coarse and Fine Particles. | Coarse particles are those with diameters greater than 2.5 μm but less than
10 μm; fine particles are those with diameters less than 2.5 μm. |  | These two fractions are usually defined in
terms of the separation diameter of a sampler.
[Note: Separation diameters other than 2.5 μm have been used.]
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-3

## Page 8
Name: “Table 1 of 1 on page 8 of 68”
Table: 8
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
5.7 | Annular. | Refers to rotating to, or forming a ring. | In the annular denuder sampler, the annular refers to
the cylinder to which coating is applied to the interior parallel planes to remove gaseous pollutants by diffusion
chemistry.
5.8 | Denuder. | To strip away the process gaseous pollutants from the gas stream.
5.9 | Equivalent Weight. | The equivalent weight, or combining weight, of a compound or ion is its formula
weight divided by the number of replaceable hydrogen atoms.
5.10 | Normal Solution. | Solution that contains a gram-equivalent weight of solute in a liter of solution.
6. Factors Affecting Denuder Efficiency
6.1 | Operation below 20% relative humidity (RH) may result in less than quantitative collection of SO2 .
Atmospheric water vapor in concentrations above 30% RH has been shown not to be an interferant for SO2
collection.
6.2 | Studies are being conducted to identify interferents, and calculations are being developed to correct the
measurements obtained by the annular denuder system for identifiable interferents. For example, the presence
of ozone (O 3 ) is known to oxidize nitrous acid (HNO2 ) to nitric acid (HNO3 ); therefore, HNO 2 |  |  | measurements
are often underestimates. Calculations have been developed to adjust for this oxidation process and to provide
more accurate estimations of HNO2 |  | concentrations in the atmosphere.
6.3 | Other studies include the possible chemical reactions (organic and inorganic) that may occur with selected
coating solutions that interfere with the accurate measurement of the chemical species of interest.
6.4 | The efficiency of impactor collection decreases when the impactor surface is loaded. |  | The average
operational time before such loading occurs has not been determined.
7. Apparatus
[Note: | This method was developed using the annular denuder system produced by University Research
Glassware, 116 S. Merritt Mill Road, Chapel Hill, NC 27516, (919) 942-2753, as a guideline. |  |  | EPA has
experience in use of this equipment during various field monitoring programs over the last several years.
Other manufacturers' equipment should work as well. Denuder systems are also available from Rupprecht
and Patashnick Co., Inc., 25 Corporate Circle, Albany, NY 12203, (518) 452-0065 and Ogawa & Company,
USA, Inc., 1230 S.E. 7th Avenue, Pompano Beach, Florida 33060, (305) 781-6223. However, modifications
to these procedures may be necessary if another commercially available sampler is selected.]
7.1 | Sampling
| 7.1.1 Elutriator and Acceleration Jet Assembly (see Figure 3). |  | Under normal sampling conditions, the
elutriator or entry tube is made of either Teflon®-coated glass or aluminum. When using glass, the accelerator
jet assembly is fixed onto the elutriator, and the internal surfaces of the entire assembly are coated with
Teflon®. | When aluminum is used, the accelerator jet assembly is removable. |  | The jet is made of Teflon® or
Page 4.2-4 |  | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 9
Name: “Table 1 of 1 on page 9 of 68”
Table: 9
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
polyethylene, and the jet support is made of aluminum. |  | Again, all internal surfaces are coated with Teflon®.
Both assemblies are available with 2, 3 and 4 mm inside diameter jets (nozzles).
7.1.2 Teflon® Impactor Support Pin and Impactor Frit Support Tools (see Figure 4). Made of either
Teflon® or polyethylene and used to aid in assembling, removing, coating and cleaning the impactor frit.
7.1.3 | Impactor Frit and Coupler Assembly (see Figure 5). | The impactor frit is 10 mm x 3 mm and is
available with a porosity range of 10-20 μm. |  | The frits should be made of porous ceramic material or fritted
stainless steel. | Before use, the impactor frit surface is coated with a Dow Corning 660 oil and toluene solution
and sits in a Teflon® seat support fixed within the coupler. |  | The coupler is made of thermoplastic and has
Teflon® clad sealing "0"-rings that are located on both sides of the seat support inside the coupler. |  | The
couplers are composed of two free moving female threads that house the support tools when assembling and
removing the impactor frit, and couple the denuders when sampling. Arrows on the metal band hold the female
threads together. | These arrows should be pointing in the direction of air flow (see Figure 1) when the ADS is
assembled.
[Note: In situations when there are substantial high concentrations of coarse particles (>2.5 μm), a Teflon®-
coated aluminum cyclone should be used in place of the acceleration jet and impactor assembly, as illustrated
in Figure 6.]
The cyclone is made of Teflon®-coated stainless steel. |  | The location of the cyclone with respect to the
denuder, heated enclosure, and meter box is illustrated in Figure 6.
7.1.4 | Annular Denuder (see Figure 7). | The original denuders (single channel denuders) consist of two
concentric glass tubes (see Figure 4). | The tubes create a 1 mm orifice, which allows the air sample to pass
through. | The inner tube is inset 25 mm from one end of the outer tube; this end is called the flow straightener
end. | The other end of the inner tube is flush with the end of the outer tube. | Both ends of the inner tube are
sealed. | In this configuration, the glass surfaces facing the orifice are etched to provide greater surface area for
the coating. | Three types of denuders are available. | One is the older version that accommodates the impactor
support pin assembly and can only be the first denuder in sequence. |  | It is available in glass with the impactor
support holder made of glass and the impactor support pin assembly made of Teflon®. The denuder is 265 mm
long with size #30 threads for coupling. | It is available with flow straighteners at both ends; however, most
denuders in use today only have one flow straightener end. |  | The second most recent denuder version, which
can be used as any denuder in sequence, is available in glass with only one flow straightener end. |  | It is 242 mm
long and has size #30 threads. | Finally, the third denuder (multi-channel denuder) design involves two inner
concentric glass tubes (1 mm separation) positioned around a solid center glass rod, as illustrated in Figure 7.
Once again, the glass surfaces are etched to provide greater surface area for the coating. |  | The inner glass tubes
and coater rod are inset 25 mm from one end of the outer Teflon®-coated stainless steel tube to serve as the
flow straightener end. Since 1992, the multi-channel denuder has been evaluated and used extensively by EPA
in place of the single channel, all glass denuder. |  | The stainless steel sheath multi-channel denuder, while more
expensive, is easier to handle during coating, extraction procedures, and in field applications. |  | Additional
channels in the denuder increases capacity of the denuder, thus minimizing breakthrough.
This procedure uses two denuders in series. The first denuder, Na 2CO3 coated, is used to trap out HNO 3
and SO 2. The Na2CO3 denuder can also trap out HCl, HNO 2, HF, and some organic acids. However, special
care needs to be taken in the preparation of the denuder coating material and extraction procedures to
quantitate HCl. | This method does not quantitate HCl, HNO2, HF, or other organic acids retained on the
Na2CO3 denuder.
The second denuder is coated with citric acid for trapping NH 3. |  | This procedure does cover the
quantitation of NH3.
All denuder types should be equipped with thermoplastic or polyethylene caps when purchased.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-5

## Page 10
Name: “Table 1 of 1 on page 10 of 68”
Table: 10
Method IO-4.2 | [10 empty] | Chapter IO-4
Acidic/Basic Constituents | [8 empty] | Atmospheric Acidic
7.1.5 | Caps for Annular Denuder (see Figure 2). |  |  |  |  | Caps are made of either polyethylene or thermoplastic
and are used in the coating and drying processes and for storage and shipment. The thermo-plastic caps include
a removable Teflon® seal plate when purchased. |  |  |  |  | Repeated reuse of these types of caps have caused some
contamination due to improper cleaning of the cap and Teflon® seal plate (i.e., fluid tends to be trapped under
the seal plate). | The polyethylene caps are not equipped with seal plates. | [6 empty] | Polyethylene caps tend to dry faster
and seal better than the thermoplastic caps.
7.1.6 | Annular Denuder Couplers (see Figure 4). |  |  |  |  | The couplers should be made of thermoplastic and
equipped with Teflon® "O"-rings that sandwich a silicone rubber ring on three sides. | [8 empty] | This design provides
elasticity for better sealing under extremely cold temperature conditions in which Teflon® does not give. | [10 empty] | Two
types of couplers are available. |  | In the older version, the couplers have removable seal rings. | [8 empty] | Problems with
denuder breakage and leakage due to improper threading of the couplers with the denuders led to the
development of a second type of coupler. |  |  |  | The new couplers are equipped with permanent seal rings that
provide more even threading and a better seal when coupled. | [6 empty] | Some couplers have built-in flow-straighteners.
The couplers are used to couple the annular denuders and for coupling the last denuder with the filter pack.
7.1.7 Filter Pack Assembly (see Figure 8). The filters are supported by stainless steel porous screens and
housed in a polyethylene filter ring housing. The Teflon® filter ring housing directly follows the Teflon® filter
housing inlet component. |  | The "nylon" filter ring housing follows the Teflon® filter ring housing and sits on a
Teflon® "O"-ring, which seals the filter ring housing components to the filter housing outlet component. | [10 empty] | The
filter housing outlet component is aluminum and accommodates a polyethylene screw sleeve that seals the filter
pack assembly. | The sleeve is available in different lengths to accommodate up to four filter ring housing units.
A stainless | steel | "Quick-Release" | plug | screws | into | the | aluminum | outlet | component | for | connecting | the
pump-timer to the filter pack assembly. It is equipped with an orange "dust cover" (male plug) upon purchase.
7.1.8 | Drying Manifold Assembly (see Figure 9). |  |  |  |  | Made of pyrex and is available to accommodate as
many as four drying denuders. |  | The denuders are attached to the manifold with back-to-back Bakalite bored
caps. | The bored caps are connected with a Teflon® connector ring. Air is pushed through an air dryer/ cleaner
bottle made of 2 1⁄2" heavy wall pyrex that contains silica gel, calcium sulfate, and activated charcoal (not
available with assembly). |  | The tubing that connects the dryer/cleaner bottle to the drying manifold should be
secured at each cap with either Teflon® washers or Teflon® washers coupled with Teflon® hose barbs.
A mixture of activated carbon/drying agent/sodium carbonate flux (~50 g) is contained in the
dryer/cleaner bottle to scrub contaminants from the supply air source.
7.1.9 | Vacuum Tubing. | Low density polyethylene tubing, 3/8" diameter for distances of less than 50 ft.,
1⁄2" diameter for distances greater than 50 ft. |  |  |  | Since this tubing is used downstream from the sampler, similar
sized tubing or pipe of any material may be substituted. | [5 empty] | The tubing must have sufficient strength to avoid
collapsing under vacuum [Fisher-Scientific, 711 Forbes Ave., Pittsburgh, PA 15219, (412-787-6322)].
7.1.10 | Tube Fitting. | Compression fittings (Swagelok®, Gyrolok® or equivalent) to connect vacuum
tubing (above) to an NPT female connector or filter holder and connect vacuum tubing to fitting on differential
flow controller. | The fittings may be constructed of any material since they are downstream of the sampler
[Fisher-Scientific, 711 Forbes Ave., Pittsburgh, PA 15219, (412-787-6322)].
7.1.11 | Annular Denuder System Sampling Case (see Figure 10). | [6 empty] | Made of a "high-impact" plastic and
insulated with polyurethane. |  | It is 4 ft long by 6" wide and 6" deep. |  |  |  |  | Two heater units, a fan blower, and an air
outlet are located in the lid of the housing. |  |  |  | Also, located on the lid are the automatic and manual control
switches and a 12-V power supply outlet for the heater and fan. | [6 empty] | The bottom of the box houses the ADS. |  |  |  | The
elutriator end of the ADS protrudes through one end of the box, while the denuders are supported in the box
by chrome plated spring clips. |  | If the Teflon®-coated aluminum cyclone is used to remove coarse particles, it
is also housed in the heated sampling box, with the elutriator end protruding through the sampling box. | [11 empty] | A
vacuum plug known as a "quick-release" coupler is linked to the filter pack of the ADS. This plug connects the
ADS to 1 1/4" | Teflon® rubber "clad" shrink tubing that exhausts the air stream to the ambient air. | [9 empty] | The box
is sledge hammer proof.
Page 4.2-6 |  | Compendium of Methods for Inorganic Air Pollutants | [8 empty] | June 1999

## Page 11
Name: “Table 1 of 1 on page 11 of 68”
Table: 11
Chapter IO-4 |  |  |  |  | Method IO-4.2
Atmospheric Acidic |  |  |  | Acidic/Basic Constituents
7.1.12 Annular Denuder Field-to-Lab Case (Optional). Annular Denuder Field-to-Lab Case (Optional).
The field-to-lab case is made of rigid plastic and insulated with polyurethane. It is made to be hand carried, not
shipped, and is used to transport four total annular denuder systems, each consisting of either three annular
denuder sections or two annular denuder sections and one denuder-impactor assembly. The systems are packed
already assembled and capped; they are ready for sampling or sample analysis. The case has a carrying handle,
a lock, and three latches and is equipped with two keys.
| 7.1.13 | Annular Denuder Shipping Case (Optional). |  | Made of formica, backed with plywood and
insulated with polyurethane. The corners are reinforced with metal. It is made to withstand shipping by truck,
UPS, and Federal Express. Each case is stackable and lockable and has a carrying handle. Seven total annular
denuder systems can be packed in the case, provided each system contains four denuders each. |  |  |  |  | The systems
can consist of either three denuders (242-mm long) and one denuder-impactor assembly (265 mm long) or four
denuders (242-mm long). |  |  | Each component of the system is packed in its own storage compartment. |  | The
personal sampler assemblies can also be placed and shipped in this case.
| 7.1.14 | Differential Flow Controller (Pump). |  | Pumps air through the sampler at a fixed rate of between
5 and 20 standard L/min (typically 10 L/min) with a precision of ±5% over the range of 25 to 250 mm Hg
vacuum.
| 7.1.15 | Dry Gas Meter (DGM). |  | Pulls 10 L of gas per revolution [Fisher-Scientific, 711 Forbes Ave.,
Pittsburgh, PA 15219, (412-787-6322)].
7.2 | Analysis
| 7.2.1 | Ion Chromatograph. | A chromatograph equipped with the appropriate anion and cation exchange
resin filled separator and suppressor columns and conductivity detector for measuring acidic (SO2 , HNO 2 and
HNO 3 ) and basic (NH3 ) ions in solution (i.e. denuder and filter extracts) [Dionex Corp., 1228 Titan Way,
Sunnyvale, CA 94086, (408-737-0700)].
| 7.2.2 Technicon Colorimeter Autoanalyzer. |  |  | Colorimetic analyzer able to detect specific ions of interest
in aqueous extracts [Technicon Industrial Systems Corp., 511 Benedict Ave., Tarrytown, NY, 10591-5097,
(800-431-1970)].
| 7.2.3 | pH Meter. | A pH or pH/ion meter with an "integral" automatic temperature compensation and
calibrated with standard buffers (pH 4 and 7), including 2 and 4 mL analysis cups (Orion and other vendors).
| 7.2.4 | Polyethylene Bottles with Polyethylene Screw Caps. |  | 50 mL and 100 mL. | Used for storage of
coating solutions, best source.
| 7.2.5 Erlenmeyer Flasks. |  | 250 mL and 2 L borosilicate glass or polyethylene flasks with calibration, best
source.
| 7.2.6 Graduated Cylinders. |  | 10 mL and 100 mL borosilicate glass or polyethylene cylinders, best source.
| 7.2.7 | Pipets. | Class A 5 mL and 10 mL borosilicate glass pipettes or automatic pipettes. |  | Calibrated "to
deliver," best source.
| 7.2.8 | Pipet Bulb. | Made of natural rubber. | Recommended to meet OSHA requirements, best source.
| 7.2.9 Micropipettes. |  | Recommended 50 μL. Calibrated "to contain," borosilicate glass micropipette, best
source.
7.2.10 Forceps. Recommended dressing forceps made of stainless steel or chrome-plated steel and without
serrations. |  | Used for handling filters (Millipore).
| 7.2.11 | Stopwatch. | Used for measuring flow rate of gas stream through DGM, best source.
| 7.2.12 | Ultrasonic | Cleaner. | Used for filter extractions and parts cleaning. | Most are temperature
controlled. |  | Temperature should be controlled during extraction at 65EC [Cole-Palmer Instrument Co., 7425
N. Oak Park Ave., Chicago, IL 60648, (980-323-4340)].
| 7.2.13 | Clean Air Hood. | Closed air hood with ammonia free air circulation. | Used for Teflon® filter
extraction for pH analysis, best source.
June 1999 |  |  | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-7

## Page 12
Name: “Table 1 of 1 on page 12 of 68”
Table: 12
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
8. Reagents and Materials
8.1 | Teflon® Filters. | Zefluor® (PTFE) membrane filters 47-mm diameter with a 2 μm pore size. |  | Only one
side is Teflon®-coated; this side should face the air stream [Gelman Sciences, 600 S. Wagner Rd., Ann Arbor,
MI, 48106, (800-521-1520)].
8.2 | Nylasorb® Filters. | Membrane filters 47-mm diameter with a 1 μm pore size. | These filters are specially
prepared and batch analyzed for low NO3 G background levels. |  |  | If other brands of nylon membrane filters are
used, they should be batch analyzed to ensure low and replicable levels of |  |  | NO3 G [Gelman Sciences, 600 S.
Wagner Rd., Ann Arbor, MI, 48106 (800-521-1520)].
8.3 | Denuder Extract Storage Vials. | 30 mL (1 oz) screw-cap polyethylene sampling vials (Nalgene or
equivalent). |  | Allow eight per sample for each sampling period, best source.
8.4 Filter Extract Storage Vials. |  | 25 mL polyethylene vials (Nalgene or equivalent). Allow two vials for each
sampling period, best source.
8.5 | IC Analysis Vials and Caps. | Available in 5 mL and 0.5 mL and are made of polypropylene. |  | The filter
caps are made of plastic and contain a Teflon® filter through which the sample is extracted for analysis. Both
the vials and filter caps should be disposable, best source.
8.6 | Labels. | Adhesive, for sample vials, best source.
8.7 | Parafilm. | Used for covering flasks and pH cups during pH analysis, best source.
8.8 | Kimwipes® and Kay-dry Towels. | Used for cleaning sampling apparatus and analysis equipment, best
source.
8.9 | Stoppers. | Cork or polyethylene, best source.
8.10 | Sodium Carbonate (Na 2 CO 3 ). | ACS reagent grade, best source.
8.11 | Sodium Chloride (NaCl). | ACS reagent grade, best source.
8.12 | Toluene. | ACS reagent grade, best source.
8.13 | Glycerol (Glycerine - CH2 OHCHOHCH2 OH). |  | ACS reagent grade, best source.
8.14 | Citric Acid [Monohydrate - HOC (CH 2 CO) OH] 2 COOH : H2 O. |  | ACS reagent grade, best source.
8.15 | Hydrogen Peroxide (H 2 O2 ). | ACS reagent grade, best source.
8.16 | Ethanol (C2 H5 OH). | ACS reagent grade, best source.
8.17 | Sulfuric Acid (H2 SO4 ). | ACS reagent grade, best source.
Page 4.2-8 |  | Compendium of Methods for Inorganic Air Pollutants |  | June 1999

## Page 13
Name: “Table 1 of 1 on page 13 of 68”
Table: 13
Chapter IO-4 | [11 empty] | Method IO-4.2
Atmospheric Acidic | [9 empty] | Acidic/Basic Constituents
8.18 | Potassium Chloride (KCl). |  |  |  | ACS reagent grade, best source.
8.19 | Perchloric Acid (HClO4 ). |  |  | ACS reagent grade (60-62%), best source.
8.20 | Distilled Deionized Water (DDW). |  |  |  |  | ASTM Type I water.
8.21 | pH Buffers. | Standard buffers 4.00 and 7.00 for internal calibration of pH meter, best source.
8.22 | Silica Gel. | ACS reagent grade (indicating type), best source.
8.23 | Sodium Bromide (NaBr). |  |  | ACS reagent grade, best source.
8.24 | Activated Charcoal. |  | ACS reagent grade, best source.
8.25 Balance. Electronic analytical with internal calibration weights and enclosed weighing chamber. Precision
of 0.1 mg [Fisher-Scientific, 711 Forbes Ave., Pittsburgh, PA, 15219, (412-787-6322)].
8.26 Gloves. | Polyethylene disposable. Used for impactor frit assembly and filter pack assembly, best source.
8.27 | Dow Corning High Temperature Vacuum Oil. | [6 empty] | Dow Corning 660 oil used for impactor frit coating
solution, best source.
8.28 | Zero Air. | A supply of compressed clean air, free from particles, oil, NO, NO2 , SO2 , HNO 3 , and HNO 2 .
The supply may be either from a commercial cylinder or generated on site, best source.
8.29 | Ion Chromatographic (IC) Eluent Solution. | [6 empty] | For extracting filters. |  | This solution should be the same
eluent as used for the ion chromatographic analysis of the filters. | [8 empty] | If the filter analysis is not to be performed
by ion chromatography, a slightly basic solution (e.g., 0.003 N NaOH or sodium carbonate/bicarbonate) should
be used to extract the Nylasorb® filter, while the Teflon® filter should be extracted with double-distilled
deionized water (DDW).
9. Preparation of Coating and Extraction Reagents
9.1 Impactor Frit Coating Solution Preparation. | [7 empty] | Weigh 1 g of silicone oil (Dow Corning high temperature
660 oil) and place in a 100 mL polyethylene storage bottle. | [7 empty] | Add 100 mL of toluene. |  |  | Mix thoroughly, close
container, and store at room temperature. | [5 empty] | (WARNING - FLAMMABLE LIQUID).
9.2 | Impactor Frit Extraction Solution Preparation. | [6 empty] | Add 100 mL of IC eluent to a clean polyethylene
storage | container. | Pipette | 5 | mL | of | ethanol | into | container. | Mix | thoroughly. | Store, | covered | at | room
temperature.
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | [9 empty] | Page 4.2-9

## Page 14
Name: “Table 1 of 1 on page 14 of 68”
Table: 14
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
9.3 | Annular Denuder Coating Solutions Preparation
[Note: Different coatings may be used depending on the chemical species of interest.]
9.3.1 Na 2 CO3 Coating Solution. Clean a 100 mL polyethylene storage vial and let dry at room temperature.
Measure 50 mL of ethanol (WARNING - TOXIC, FLAMMABLE LIQUID) with a graduated cylinder and pour
into vial. | Measure 50 mL of DDW with a graduated cylinder and add to vial. |  |  | Weigh 1 g of glycerol and add
to DDW. Weigh 1 g of Na2 CO3 | and add to vial. |  | Mix thoroughly, solution may fizz; wait for fizzing to stop
before sealing vial. | Store at room temperature.
9.3.2 | Citric Acid Coating Solution. | Clean a 100 mL polyethylene storage vial and let dry at room
temperature. | Measure 50 mL of ethanol (WARNING - TOXIC, FLAMMABLE LIQUID) with a graduated
cylinder and pour into vial. | Weigh 0.5 g | of | citric acid and add to a vial. | Add enough glycerol to vial to make
it a 1% solution (0.5 mL). | Mix thoroughly; store, covered at room temperature.
10. Elutriator and Acceleration Jet (Inlet) Assembly
[Note: The all-glass configuration is shown in Figure 6A.]
10.1 The internal walls of the elutriator and jet assembly are coated with Teflon® to prevent losses of reactive
species (SO 2 , HNO 3 , NH 3 ) during sampling. The elutriator prevents water and large particles from entering the
inlet and thus extends the life of the impaction surface located immediately downstream of this assembly.
10.2 | An aluminum version of this inlet is shown in Figure 3b. |  | All inner surfaces of the aluminum unit are
Teflon®-coated. | The main difference between the all glass and the aluminum inlet is the jet component of the
aluminum inlet is replaceable, as shown in Figure 3b. |  |  | The jet component is made of either Teflon® or
polyethylene and is available in various diameters as needed to accommodate selected sample flow rates. |  |  |  | The
jet may be replaced using the tool shown in Figure 3b. |  |  | The jet diameter for a sample flow rate of 10 L/min is
3.33 mm. | At this flow rate, the inlet has an aerodynamic cutpoint of 2.5 μm. |  |  | If a different flow rate is used,
the jet diameter must be changed to retain an aerodynamic cutpoint to 2.5 μm. |  |  |  | Table 2 contains the jet
diameters and Reynolds number to maintain an aerodynamic of 2.5 μm cutpoint at different flow rates between
1 and 20 L/min.
[Note: | If the sampling area has substantial concentrations of coarse particles (>2.5 μm), the user may
replace the acceleration jet and impactor assembly with the Teflon®-coated aluminum cyclone.]
11. Impactor Frit Preparation and Installation
11.1 | Impactor Frit Installation
11.1.1 Impactor-Coupler. The impactor-coupler assembly (shown in Figure 12) is composed of two parts:
the replaceable impactor frit and the coupler-impactor housing seat. |  |  | The impactor surface is a porous ceramic
or porous stainless steel frit, 10 mm x 3 mm. |  |  | This frit is inserted into the coupler-impactor housing using the
tools shown in Figure 12. | The in-tool must be completely screwed in behind the impactor seat before the frit
is pressed into place. | Press the impactor frit gently but firmly into the seat of the impactor housing with your
Page 4.2-10 | Compendium of Methods for Inorganic Air Pollutants |  |  | June 1999

## Page 15
Name: “Table 1 of 1 on page 15 of 68”
Table: 15
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
clean gloved finger. | The impactor should fit into the housing so that it does not protrude above the seat. |  | The
impactor frit has a slight bevel. |  | The narrow surface should be inserted into the impactor seat.
11.1.2 | Impactor-Denuder. | The impactor-denuder assembly (shown in Figure 4) is of three parts: | the
replaceable impactor frit, the impactor seat support pin, and the annular denuder impactor-pin support. |  |  | The
impactor frit is the same as described in Section 11.1.1 and is inserted into the impactor seat support pin. |  |  | The
impactor support pin can be hand-held while inserting the frit or it can be placed upright into the stainless steel
frit holder #3 (see Figure 11). | Press the support pin into the denuder pin support. |  | The pin is grooved and has
a viton "O"-ring to keep the pin snug in the denuder support during cold weather use (Teflon® tends to shrink
at low temperatures). | The support pin is removed by using the removal tool shown in Figure 4.
11.2 | Impactor Frit Preparation
With the impactor frit in the impactor seat of either the coupler (see Figure 12) or the Teflon® impactor seat
support pin that fits into the first denuder (see Figure 4), pipette 50 μL of the toluene-660 oil coating solution
onto the impactor frit surface and allow to dry at room temperature. Cap both sides of the coupler impactor
or denuder-impactor until use.
12. Filter Pack Preparation and Assembly
[Note: | Any number of filters can be used depending on the target species of interest. |  | The configuration
referred to in this section does not collect NH +4 |  | .]
12.1 With clean gloves, disassemble the filter pack (see Figure 8) by unscrewing the large outer Teflon® collar
(sleeve) from the aluminum filter housing outlet component.
[Note: Remove the red Bakelite® cap first. Lay the pieces out on clean Kimwipes®.]
12.2 | Lay a clean Teflon® filter ring housing, with its large opening face up, on a clean Kimwipe®. |  | Place a
clean stainless steel screen in the filter ring housing.
12.3 Using clean filter forceps, place a Nylasorb® nylon filter on the screen. Insert a second filter ring housing
on top of the nylon filter with its large opening face up. |  |  | This design forms a "sandwich" with the nylon filter
held between the two filter ring housings.
12.4 | Place another clean screen on the second filter ring housing. |  | Using clean filter forceps, place a Teflon®
filter on the screen.
[Note: If a Teflasorb® Teflon® filter is used, be sure to place the Teflon®-coated side, not the webbed side,
toward the air stream. If the webbed side is facing the air stream, SO 4= extraction from the filters may be
inefficient.]
12.5 Place the Teflon® filter housing inlet component (see Figure 11) on top of the Teflon® filter, which forms
another "sandwich" with the Teflon® filter held between the second filter ring housing and the housing inlet
component. The housing inlet component connects the filter pack assembly to the last annular denuder through
a thermoplastic coupler. | Be careful not to twist the filterpack components, or damage will occur to the filters.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-11

## Page 16
Name: “Table 1 of 1 on page 16 of 68”
Table: 16
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
12.6 | Lay the aluminum filter housing outlet component, with its large opening face up, on a clean Kimwipe®.
12.7 | Insert the filter ring sandwiches (prepared in Sections 12.1 through 12.5) with the filter housing inlet
component extending upward, in the aluminum filter base. |  | Place the large outer Teflon® sleeve over the filter
sandwich and screw onto the aluminum filter base. |  | DO NOT OVERTIGHTEN! | DO NOT TWIST FILTER
PACK COMPONENTS!
12.8 | Install the "Quick-Release" plug into the filter outlet component. | DO NOT OVERTIGHTEN!
12.9 | Install the red Bakalite® cap onto the filter inlet component and the orange dust cover onto the
Quick-Release plug until ready to attach denuders.
13. Annular Denuder System Preparation
All new annular denuder parts obtained from suppliers should be cleaned with a dilute soap solution. The
parts should then be thoroughly rinsed in DDW and allowed to dry at room temperature.
13.1 | Annular Denuder Coating Procedure
[Note: If the first denuder holds the impactor, a blank Teflon® impactor support pin should be installed in
the pin support holder before the coating procedure.]
13.1.1 | Cap the end of the denuder that has the inner tube flush to the outer tube and set denuder upright
on the capped end. | For the denuders with flow-straighteners at both ends, either end can be capped. |  | Measure
10 mL of the appropriate coating solution into a graduated cylinder. Pipette the 10 mL into the flow-straightener
end of the upright capped annular denuder.
13.1.2 | Cap the open end of the denuder, and holding horizontally, rotate the denuder to distribute the
coating solution evenly.
13.1.3 | Remove cap from flow-straightener end of denuder and decant excess coating solution into a clean
denuder extract storage bottle labeled "denuder blank." |  | Bottle label should include denuder number, coating
solution and date.
13.1.4 | Repeat this procedure with each denuder; label the denuders and bottles appropriately.
13.2 | Annular Denuder Drying Procedure (see Figure 9)
[Note: As denuders dry, they change from translucent to a frosted appearance. Denuders are dry when they
become uniformly frosted.]
13.2.1 Adjust drying train and manifold clean air flow to 2 to 3 L/min. Close toggle valve controlling clean
air flow through manifold before attaching denuders.
13.2.2 Attach flow-straightener end to drying manifold port at the back-to-back bored caps (see Figure 9).
13.2.3 | Open toggle valve and allow clean air to flow through the tube for several minutes.
13.2.4 | Close toggle valve and reverse ends of tubes attached to manifold.
13.2.5 | When an even frosted appearance is achieved, remove tubes from manifold, cap both ends with
clean caps, and store until ready for use. | Turn off air to drying manifold.
Page 4.2-12 | Compendium of Methods for Inorganic Air Pollutants |  | June 1999

## Page 17
Name: “Table 1 of 1 on page 17 of 68”
Table: 17
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
13.3 | Annular Denuder System (ADS) Assembly
[Note: Described herein is an annular denuder system consisting of four denuders in series. Any number of
denuders can be used at the operators' discretion. The denuders should be assembled in such a way that the
flow-straightener end always follows the flush end of the previous denuder, except, if denuders with flow-
straighteners at both ends are used. |  | This type of assembly allows laminar flow conditions to be restored
quickly.]
13.3.1 | Lay the ADS pieces on a clean surface (i.e., Kimwipes®).
13.3.2 | Remove the end caps from the first denuder. | Denuder 1 is coated with Na2 CO3 | and holds the
impactor frit pin support. | Remove the blank impactor support pin. | Gently insert the impactor support pin and
coated frit assembly into the denuder-pin support. |  | If the first denuder does not hold the impactor pin-support,
attach the impactor frit seat equipped coupler assembly to the flow-straightener end of the first denuder.
[Note: DO NOT TIGHTEN! Do not tighten during the following procedure until Section 13.4.8 is reached.]
13.3.3 | Attach a thermoplastic coupler to the opposite denuder end. | Place a Teflon® clad "O"-ring inside
the coupler, if needed.
13.3.4 | Remove the end caps from the second denuder (citric acid coated). | Attach the end with the
flow-straightener section to the denuder-coupler assembly.
13.3.5 | Attach a thermoplastic coupler to the opposite denuder end. | Place a Teflon® clad "O"-ring inside
the coupler, if needed.
13.3.6 | Attach the filter pack inlet to the second denuder coupler assembly.
13.3.7 | Attach the elutriator-acceleration jet assembly to the first denuder-coupler assembly. Tighten very
gently--DO NOT OVERTIGHTEN!
13.3.8 | Tighten the remaining couplers very gently - DO NOT OVERTIGHTEN!
13.3.9 | Cap elutriator with orange dust cover until use.
[Note: | When collecting and measuring gaseous HNO 3, SO 2, and particulate NO 3- and SO 4= , NH 3 must be
taken out of the gas stream prior to the air stream entering the filter pack. |  | Otherwise, reaction of the
unneutralized sulfate will result. If ammonia (NH 3) and/or H + measurements are not to be analyzed for, then
the use of a citric acid coated denuder is not important. |  | However, with the removal of NH 3, some nitrate
collected on the Teflon® filter will evaporate and be found on the nylon filter.]
13.4 | Laboratory Leak-Check of ADS
[Note: CAUTION--Do not subject the system to sudden pressure changes or filters may tear.]
13.4.1 | Remove the orange dust cap from the impactor opening. | Attach the "Quick-Release" to a pump
module. | Turn on the pump. | Be certain that flow through the ADS occurs by checking the rotameter.
13.4.2 | Briefly cap the elutriator with the orange dust cap. | The flow as indicated on the rotameter should
drop to zero if no leaks exist.
13.4.3 | Disconnect the pump from the ADS at the "Quick-Release" plug. | Cap the "Quick-Release" plug
with an orange dust cover. Turn off the pump. REMEMBER--never overtighten joints or breakage will result.
If the joints can not be sealed with gentle tightening, the Teflon® "O"-rings are worn or defective and must be
replaced.
13.4.4 | Place the assembled sampler in its field-to-lab carrying case for transport to the field.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-13

## Page 18
Name: “Table 1 of 1 on page 18 of 68”
Table: 18
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
[Note: The ADS joints should be loosened slightly when extreme temperature changes are incurred during
transportation. |  | This precaution will prevent unnecessary breakage or distortion of the ADS components.
Remember to allow the system to adjust to the indoor air temperature before tightening the joints and checking
for leaks.]
14. Sampling
14.1 | Placement of Denuder System
14.1.1 | The placement of the annular denuder system must conform to a consistent set of criteria and
guidance to ensure data comparability and compatibility. |  | A detailed set of monitor siting criteria for ambient
air monitoring and meteorological programs is given in the EPA document Ambient Monitoring Guidelines for
Prevention of Significant Deterioration (PSD), EPA-450/4-87-007, U. S. Environmental Protection Agency
Office of Air Quality Planning and Standards, Research Triangle Park, NC 27711, May 1987.
The site must be away from localized sources of ammonia, such as composting and livestocking
operations, landfills, sewage treatment plants, fertilizer plants and storage facilities, and recently plowed
fertilized fields because aerosol acidity is subject to rapid neutralization by ambient bases.
14.1.2 | A summary of key factors that should be considered as part of the placement of an air quality
monitoring station containing an ADS are:
• Vertical placement above ground;
• Horizontal spacing from obstructions and obstacles;
• Unrestricted air flow; and
• Adequate spacing from roads.
The ADS sampler is mounted on a supported mast pole or tripod. The ADS inlet should be located 2-3
m above ground level. Placing the inlet closer to ground level should be considered only if the surface is flat
and man-made (i.e., not unpaved dirt).
14.1.3 A summary of key criteria associated with these siting factors for air monitoring stations is included
in Table 3.0. | The information included in the table should be used to the extent possible as part of the
monitoring network design to ensure that the monitoring program provides representative and unbiased data.
However, site-specific constraints could make it very difficult to meet all criteria. For example, wooded areas
around a site would make the siting very difficult. |  | The use of the information in Table 3.0, coupled with a
balanced evaluation by an experienced air quality and meteorology specialist, is highly recommended.
14.1.4 In general, for a site with no major obstruction and obstacles, the air sampler intake should be about
2-3 m aboveground. For a site with nearby roadways, however, intake placement should take into account the
effects of road dust re-entrainment and vehicular emissions. |  | In fact, a linear relationship should be established
between the horizontal distance of the sampler intake from the roadway and the aboveground elevation of that
intake. | For any roadway accommodating more than 3,000 vehicles per day, the intake should be between 5
and 25 m from the edge of the nearest traffic lane. |  | It should also be 15 m aboveground for a distance of 5 m
from the nearest traffic lane and 2 m aboveground for a distance of 25 m from the nearest lane. For a roadway
supporting less than 3,000 vehicles per day, the intake should be placed at a distance greater than 5 m from the
edge of the nearest traffic lane and at a height of 2-15 m aboveground.
14.2 | Start-Up
14.2.1 Remove the ADS from its field-to-lab carrying case and load into the field sampling box. |  |  | The ADS
field sampling box is insulated with polyurethane which is configured to hold the ADS without allowing
movement. | Chrome plated spring clips hold the denuders in place. | Automatic and manual control switches
Page 4.2-14 |  | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 19
Name: “Table 1 of 1 on page 19 of 68”
Table: 19
Chapter IO-4 | [12 empty] | Method IO-4.2
Atmospheric Acidic | [11 empty] | Acidic/Basic Constituents
allow the sampling box to control the temperature of the ADS. | [7 empty] | The automatic switch should be used when the
ADS is not in use and when the ADS is sampling for extended periods of time without constant supervision to
prevent low temperature or sudden pressure change exposure of the ADS (these types of exposure can cause
leaks to occur, condensation, or the filters to tear). |  |  |  |  | When sampling, the ADS should be kept 1EC above the
ambient temperature to prevent condensation. The sampling box has two connections with the pump timer: the
plastic | suction | hose | connected | with | "Quick-Release" |  | couplers |  |  | and | the | 12-V | power | cord | with | a
"Quick-Disconnect" coupler. |  |  | The power cord remains connected, and the suction hose is disconnected from
the box each time the unit is opened. |  |  |  | Inside the box, the hose is connected to the top of the filter pack with a
"Quick-Release" coupler. |  |  | During sampling the sample box is kept securely closed (see Figure 2).
14.2.2 | Allow the pump to warm up for 20-30 min prior to testing so the pump will provide steady flow
during testing.
14.2.3 To check the Heat/Cool cycles, flip one switch from "AUTO" to "MANUAL" and the other between
"COOL" and "HEAT." |  |  | Check to insure that the fan and heater work, respectively.
14.2.4 With the elutriator still capped, turn on the pump with the switch on the timer. The rotameter should
indicate zero flow. |  | If there is a flow, the assembly pieces need to be recoupled. | [9 empty] | Run leak check for 5-10 s.
Turn off pump and remove elutriator cap. |  |  |  |  | Record leak rate on Field Test Data Sheet (see Figure 13).
14.2.5 | Attach DGM output to elutriator inlet. |  |  |  | Turn on pump. |  |  |  | Record start time on Field Test Data Sheet
(see Figure 13). |  | Using a stopwatch, record the time for 20.0 L to pass through the DGM. | [10 empty] | Record the DGM
temperature and the absolute pressure of the DGM.
14.2.6 | Calculate the flow rate as follows:
|  |  |  | Q |  | /P | )(T | /T | )(F | )
|  |  |  | std | bar | std | std | m | c
where:
| Q std = | flow rate corrected to standard conditions, 25EC and 760 mm Hg, L/min.
| V = | volume of gas pulled through denuder system, L.
| t = | time required to pull gas through denuder system, minutes.
| Pbar = | barometric pressure, mm Hg.
| Pstd = | standard barometric pressure, 760 mm Hg.
| Tstd = | standard temperature, 298EK.
| Tm = | temperature of dry gas meter, EK(=EC + 273).
| F c = | dry gas meter correction factor, dimensionless.
14.2.7 | If the calculated flow rate is not between 9.5 and 10.5 L/min, readjust the flow rate and repeat
Sections 14.2.4 and 14.2.5 until the rate is in the above range. | [7 empty] | Preliminary studies should be conducted to
obtain an estimate of the concentrations of the species of interest.
14.2.8 | Record the flow rate on Field Test Data Sheet.
14.2.9 Remove DGM connection tubing from elutriator inlet. Pump should remain running so that sampling
continues. Higher flow rates may be used for shorter sampling periods. Concentration of the species of interest
in indoor air and the configuration of the sampling equipment, determine the appropriate flow rates. | [13 empty] | Sampling
at 10 L/min, requires a sampling time of 24 h to collect pollutant concentrations between 0.02 and 0.83 μg/m3 .
June 1999 |  |  | Compendium of Methods for Inorganic Air Pollutants | [10 empty] | Page 4.2-15

## Page 20
Name: “Table 1 of 1 on page 20 of 68”
Table: 20
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
14.3 | Sample Shutdown
14.3.1 | Attach DGM connection tubing elutriator inlet with pump still running. |  |  | Measure flow rate as in
Sections 14.2.5 and 14.2.6. |  | Record flow time, temperature, and pressure on Field Test Data Sheet (see
Figure 13).
14.3.2 | Turn off pump. | Record time and elapsed time meter reading on log sheet. |  | Remove DGM
connection tubing from elutriator inlet. Remove ADS from the sampling box, cap the ends, and place the ADS
in field-to-lab carrying case for transport to lab. Be careful not to stress the ADS during the transfer or breakage
will result.
Caution: | When the ADS is brought from a cold field sampling location to a warm laboratory, loosen the
denuder couplings to prevent thermal expansion from breaking the denuders.
14.4 | Corrective Action for Leak Test Failure
[Note: These steps should be followed when failure occurs during testing at the laboratory before transport
to the field and in the field before testing.]
14.4.1 | Sampler Leaks. | Note the problem on the Field Test Data Sheet. |  | Check assembly of ADS
components. | Replace gaskets. | Check for proper seating of denuder surfaces. | Replace any defective parts.
14.4.2 | Cracked or Chipped Denuders or Elutriator Assemblies. |  | Note problem on Field Test Data
Sheet. | Discard defective pieces. | Do not try to extract cracked pieces. | WARNING--USE CAUTION WHEN
DISASSEMBLING CRACKED GLASSWARE. |  | Pieces may shatter and cause severe cuts. |  | Wear protective
clothing.
14.4.3 | Contaminated Blank Solutions. | Note problem on Field Test Data Sheet. |  | Follow parts-cleaning
procedures closely. | Examine the sampler preparation area for possible sources of contamination and remove
source. | Check DDW being used in the solution preparations and extractions. Fill a clean 25 mL polyethylene
extraction bottle with the DDW used in solution preparation and extraction; send to lab for analysis. |  |  |  | If
contaminated, correct deionization system.
14.4.4 Flow Rate Disagreement. |  | Note problem on Field Test Data Sheet. | Check vacuum gauge on flow
module. If a high vacuum exists, the sampler has become blocked. This blockage may be due to dust or smoke
particles clogging the filters or to obstructions in the system or tubing. |  |  | Check flow module. | Repair as needed.
14.4.5 | Inadequate Flow Rate. | Note problem on Field Test Data Sheet. |  | Check rotameter on flow
controller. | If adequate flow is shown here, a leak exists between the controller and the DGM. |  |  | If no flow is
shown on rotameter, check vacuum gauge on controller. |  |  | If no vacuum exists, pump needs repair. | If a high
vacuum is shown, an obstruction exists in the system. |  |  | Check to see that the paper filter dividers were not
accidentally installed with |  | the filters in the filter pack. | Check tubing for kinks.
[Note: Typically the pressure drop across the filters should be approximately 1" Hg at 10 L/min flow rate
at sea level. This pressure drop can vary from 1-10 L/min depending on elevation.]
15. ADS Disassembly
15.1 | Remove the ADS from the field-to-lab carrying case using both hands. |  | To prevent stress, hold the ADS
by its ends.
Caution: Do not stress the ADS while removing it from the case.
15.2 | Decouple the elutriator-jet assembly from the first denuder-impactor-coupler assembly.
Page 4.2-16 | Compendium of Methods for Inorganic Air Pollutants |  |  | June 1999

## Page 21
Name: “Table 1 of 1 on page 21 of 68”
Table: 21
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
15.3 | When using the denuder-impactor, the frit-pin must be removed from the support in the denuder before
removing the frit from the pin. |  | The frit is then extracted from the pin using pin tool #3 and the frit extraction
tool (see Figure 12). When using the impactor-coupler assembly, the frit is removed from the coupler seat using
pin tool #3 and the "out" frit removal tool (see Figure 12). |  | Put frit in covered dish and set aside for future
chemical extraction (optional).
[Note: This method does not require the analysis of the particles impacted on the frit.]
15.4 | Remove the denuders from the couplers and cover each end of the denuders with clean end caps until
extraction.
15.5 | Label a clean 100 mL polyethylene bottle with the sampler ID number and filter type (i.e., Teflon® or
Nylasorb®, as appropriate) for each of the filters.
15.6 | Disassemble the filter pack in a clean, ammonia-free air hood. | Clean all hood surfaces and utensils with
ethanol. Wearing clean gloves and using clean filter forceps, remove the filters and place each in its storage
(protective) bottle, with the exposed filter surface facing downward, until extraction.
[Note: Place the filters in the properly labeled bottles.]
16. Extraction Procedures
[Special precaution: | Samples should be analyzed as soon after collection as possible. |  | The solutions and
extraction procedures must be prepared and performed on the day of pH analysis.
Extraction must take place in a clean, ammonia-free air hood. The extracts must be processed in the order
in which they will be analyzed, so that each sample will have a similar time interval between extraction and
analysis. Denuder extracts and filters should be stored in the refrigerator until just prior to analysis. Samples
stored longer than 30 days tend to degrade due to bacteria growth and/or losses to the walls of the extraction
vessel. | Extraction volumes may vary depending upon sensitivity requirements and expected ambient
concentrations.
16.1 | Impactor Frit Coating Extraction (Optional)
16.1.1 | Place the impactor (which was removed before denuder extraction) into a small extraction bottle.
16.1.2 | Label the bottle appropriately. | Pipet 10 mL of impactor extraction solution (See Section 9.2) into
the bottle. | The solution must cover the surface of the impactor frit.
16.1.3 | Close the extraction bottle and place in an ultrasonic bath for 30 min.
16.2 | Denuder Extractions
[Note: If the denuder was the first denuder and equipped with the impactor frit-pin support, insert a clean
Teflon® impactor frit-pin without frit in place. |  | Then extract as described below. | This procedure is to be
followed for each denuder.]
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-17

## Page 22
Name: “Table 1 of 1 on page 22 of 68”
Table: 22
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
16.2.1 | Cap one end of each of the denuders (i.e., Na 2 CO3 |  | and citric acid denuders). | Add 5 mL of DDW
with a pipet. | Cap other ends of the denuders.
16.2.2 | Rotate the denuders to wet all surfaces thoroughly with the water. |  |  | Remove the caps and pour the
liquid from each of the two denuders into separate clean 25 mL polyethylene extraction bottles.
16.2.3 | Repeat this procedure with a second 5 mL of DDW extract (total extract volume is 10 mL, which
is placed into the extraction bottles).
16.2.4 Add 1 mL of hydrogen peroxide to oxidize the sulfite (SOG3 ) to sulfate (SO=4) in the extraction bottles
containing the Na2 CO3 |  | coating.
16.2.5 | Replace the extraction bottle caps and label the bottles with the sampler ID number and denuder
number and type (as appropriate).
16.3 | Filter Extraction
16.3.1 | Teflon® Filter Extraction [for pH analysis followed by SO=4, NOGG3 , and/or NH +4 analyses)
[Note: Teflon® is not wet by water; therefore, the filter will float on top of aqueous solutions. Solutions and
extraction procedures must be prepared and performed on the day of pH analysis. Extraction of the filters
must take place in a clean, ammonia-free, air hood. The filters must be processed in the order in which they
will be analyzed so that each sample will have a similar time interval between extraction and analysis.]
| 16.3.1.1 | Samples should be analyzed as soon after collection as possible. The solutions and extraction
procedures must be prepared and performed on the day of pH analysis. |  |  |  | Keep samples in a refrigerator until
extracted and analyzed.
| 16.3.1.2 | Samples should not be extracted until the day of analysis; however, if samples are extracted
and it is not possible to analyze them that day, they should be refrigerated. Allow the samples to return to room
temperature before analysis.
| 16.3.1.3 | The same extract solution (ES) must be used for the samples to be |  | analyzed, the working
standards, and the EA solution. |  |  | Also, the same batch of alcohol must be used to prepare the EA solution and
the working standards and to extract the Teflon® filters.
| 16.3.1.4 | Handling the exposed Teflon® filter requires protection from contamination with NH3 , which
may rapidly neutralize aerosol acidity on the filter and bias the sample results. |  |  |  | To ensure ammonia-free air
occupies the glove-box, a positive pressure is maintained by blowing air through a PVC tube (4" O.D.) filled
with glass-wool dosed with citric acid before entering the manifold which enables uniform distribution of air
from top of the glove-box. |  | Flow the ammonia-free air for 5 minutes before retrieving the filter. |  | Place a citric
acid soaked filter paper on the bottom to deplete ammonia when unused. |  |  |  | Disassemble the filter assembly in
the clean, ammonia-free glove-box. |  |  | Clean all glove-box surfaces and utensils with ethanol.
| 16.3.1.5 | Allow the hood to be flushed with ammonia-free air for at least 5 min before filter extraction.
All of the hood surfaces and extraction utensils must be cleaned with a Kimwipe® moistened with ethanol.
| 16.3.1.6 | Pipet 6 mL of 0.0001 N perchloric acid (HClO 4 ) solution into the appropriately labeled
extraction vials (10 mL).
[Note: Perchloric acid is used because it inhibits CO 2 from dissolving into the solution and keeps the organic
compounds in solution from dissociating. Both these activities can change the ionic strength of solution.]
| 16.3.1.7 | Wearing clean gloves and using clean filter forceps, place the Teflon® filter in the extraction
vial with the exposed filter surface facing downward. |  |  | Cap tightly. | Store at 4EC in the dark until ready for
analysis.
Page 4.2-18 |  | Compendium of Methods for Inorganic Air Pollutants |  | June 1999

## Page 23
Name: “Table 1 of 1 on page 23 of 68”
Table: 23
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
16.3.1.8 | When ready for analysis, the filter must be prepared (within the air hood) in the following
manner: | Using forceps and gloved hands, lift the filter from the extraction vial. | Let the excess solution drain
off into the vial. | Holding the filter over the extraction vial, and using an automatic pipet, apply 200 FL ± 5 FL
of ethanol to the filter. Add the ethanol slowly to ensure that all portions of the membrane are wet with ethanol.
Immerse the filter in the aqueous solution once again. |  | Tap the forceps against the inside of the vial to remove
liquid. | Tightly replace cap. | Put in ultrasonic bath for 15 min total, rotating the rack 90E every 5 min.
[Note: | Perchloric acid is used in place of potassium chloride, initially, to prevent interference in the
measurements of cations and anions by ion chromatography. |  | Potassium chloride must be added to the
portions of the sample extract that are used for pH analysis (the purpose of the salt, final concentration
0.04 M, is to increase the ionic strength and thus to reduce the time for equilibrium of the pH electrode used
for measurement). Use the same bottle (freshly opened) of ethanol to extract the Teflon® filters that are used
to prepare sulfuric acid standards.]
16.3.1.9 When ready for pH analysis, the extracts are prepared in the order of pH measurement. Inside
the air hood, remove the caps from 10 mL extraction vials. |  | Wipe off any drops which may leak onto the
outside of the cup.
16.3.1.10 | Using gloved hands and a 1 mL automatic pipet, transfer 1 mL of the extract to each of two
correspondingly labeled 2 mL cups.
[Note: The first 2 mL cup for each extract has the same I.D.# as the 10 mL cup, and the second 2 mL cup
has the same I.D.# with a hyphen (-). This system also is used with the working standards.]
16.3.1.11 | After transferring the extracts to the 2 mL cups, recap the 10 mL extraction vial. | Store the
10 mL vials at 4EC in a refrigerator pending analysis by IC, and for NH4+ analysis by Technicon auto analyzer.
16.3.2 | Nylon Filter Extraction.
16.3.2.1 | Pipet 10 mL of IC eluent into the appropriately labeled filter vial or bottle with caps.
[Note: Be sure that the filter lies flat on the bottom of the bottle and that all of the filter is covered by the
extraction solution.]
16.3.2.2 | Place the nylon filter in the extraction vial.
16.3.2.3 | Replace the bottle's cap and put in an ultrasonic bath for 30 min.
16.3.2.4 | Store the bottles in a clean (i.e., pollutant free) refrigerator at 4EC until analysis.
17. Ion Chromatography Analysis
[Note: | The analytical procedure described here is | not the only appropriate procedure available for
quantifying the analytes of interest. An automated system does not have to be used. This particular analytical
procedure was chosen because it is presently used by EPA. Modifications to this procedure may be required
depending on the intended use of the data; however, any modifications made must be justified to obtain
comparable data quality.]
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-19

## Page 24
Name: “Table 1 of 1 on page 24 of 68”
Table: 24
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
17.1 | Standards Preparation
Special Precaution: |  | Storage of these solutions should be no longer than one week. |  | All of the working
standard solutions are used to calibrate the IC and are made from reagent grade stock. The crystals are dried
overnight in covered petri dishes at 110EC in a vacuum oven prior to preparing the standard solutions. Any
yellowish discoloration of the dried crystals indicates decomposition and crystals should be discarded.
17.1.1 | Sodium Sulfate Stock Solution.
| 17.1.1.1 | In a clean, calibrated, 1 L flask, add 500 mL of DDW.
| 17.1.1.2 | On weighing paper, weigh out enough reagent (Na2 SO4 ) to make the solution 2000 ppm
concentration. |  | The target weight is 2.958 ± 0.001 g. | Record the gross weight.
17.1.1.3 Add the reagent crystals to the 500 mL of DDW. Reweigh weighing paper and subtract weight
from the gross weight. |  | The difference is the actual net weight.
| 17.1.1.4 Dilute to volume with DDW. | Mix well and cover with parafilm.
17.1.2 | Sodium Nitrate Stock Solution.
| 17.1.2.1 | In a clean, calibrated, 1 L flask, add 500 mL of DDW.
| 17.1.2.2 | On weighing paper, weigh out enough reagent (NaNO3 ) to make the solution 2000 ppm
concentration. |  | The target weight is 2.742 ± 0.001 g. | Record the gross weight.
| 17.1.2.3 | Follow Sections 17.1.1.3 and 17.1.1.4.
17.1.3 | Standard working solutions (100 ppm SO 4= and 100 ppm NO3- ). |  |  | The working solutions are
made as follows: Add 10 mL each of the two stock solutions (SO4= and NO 3 G) to a 200 mL volumetric flask
and dilute to the mark with DDW. |  | Subsequent dilutions are carried out using a 10 mL volumetric pipet and
appropriate flasks. |  | Standards of 20, 10, 5 and 1 ppm | SO4= and NO 3 G are prepared to calibrate the IC.
17.2 | Reagent Preparation
[Note: Storage of these reagents should be no longer than 1 week.]
17.2.1 | Anion eluent. | The anion eluent is a solution of 1.8 mM Na 2 CO3 |  | and 1.7 mM NaHCO 3 . | A
concentrated solution can be prepared and diluted as needed.
[Note: See Anion Storage Solution]
| 17.2.1.1 Concentrated Na 2 CO 3 | solution (0.36 M). | Weigh out 38.156 g of Na2 CO3 (MW = 105.99).
Dissolve into 1 L of DDW. |  | Store in refrigerator until ready to dilute.
| 17.2.1.2 Concentrated NaHCO 3 solution (0.34 M). |  | Weigh out 28.564 g of NaHCO3 (MW = 84.01).
Dissolve into 1 L of DDW. |  | Store in refrigerator until ready to dilute.
| 17.2.1.3 Dilution of stock solutions. | Bring both solutions to room temperature. |  | Accurately pipet 10
mL of each solution into a 2000 mL volumetric flask which has been partially filled with DDW. |  |  |  | Bring to the
mark with DDW (1:200 dilution).
17.2.2 | Anion regenerant. | The regenerant is a 0.025 N H 2 SO4 solution. | VERY CAREFULLY dispense
2.8 mL of concentrated Ultrex sulfuric acid (36 N) into a graduated cylinder. |  |  |  | Partially fill the regenerant
reservoir with DDW (3 L). |  | Slowly add the acid to the regenerant reservoir. | Bring to the mark with DDW
(4 L).
[Note: Protective clothing and eye protection should be worn.]
Page 4.2-20 |  | Compendium of Methods for Inorganic Air Pollutants |  | June 1999

## Page 25
Name: “Table 1 of 1 on page 25 of 68”
Table: 25
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
17.2.3 | Cation eluent. | There are two cation eluents that are used for the analysis of monovalent and
divalent cations. |  | The strong cation eluent is: 48 mM HCl, 4 mM DAP.HCl (DAP = Diaminoproprionic acid),
4 mM Histidine.HCl. |  | The weak eluent consists of 12 mM HCl, 0.25 mM DAP.HCl, 0.25 mM Histidine.HCl.
| 17.2.3.1 Strong cation eluent. |  | Weigh 0.560 g DAP and 0.840 g histidine into a 1 L volumetric flask.
Add 48 mL of 1 M HCl (Ultrex) to the flask. |  |  | Bring the eluent to the final volume by bringing to the mark with
DDW. | Mix thoroughly to dissolve.
| 17.2.3.2 | Weak cation eluent. | Place 63 mL of the strong cation eluent in a 1 L flask. Add 9 mL of 1
M HCl to the flask. |  | Bring the eluent to the final volume by bringing to the mark with DDW. Mix thoroughly
to dissolve.
17.2.4 | Cation regenerant. | The cation regenerant consists of 100 mM tetrabutyl-ammonium hydroxide
(TBAOH). Place the TBAOH container into a warm water bath to dissolve any crystals that may have formed.
Measure 266.7 mL of the TBAOH (stock reagent is supplied as 1.5 M, 40% in water) into a graduated cylinder.
Add the TBAOH to 4 L of DDW.
17.2.5 | Anion storage solution. | Since the anion columns contain carbonates from the eluent, protection
must be taken against microorganisms that will live on this food source and clog up the columns. If the columns
are not being used for long periods of time (>2 weeks), a storage solution of 0.1 M NaOH should be pumped
into them.
17.3 | Sample Preparation
17.3.1 | Mark the auto sampler vials with the appropriate identification numbers. |  | Place the vials in an IC
autosampler tray.
17.3.2 | Using clean, calibrated 0.5 mL pipets, transfer the denuder extract and the remainder of the filter
extracts from the extraction vials to a clean disposable 0.5 mL IC autosampler (polyethylene) vial. |  |  | Fill the
autosampler vial to the line on the side.
[Note: If refrigerated, the contents of the 10 mL extraction vial must be vortex-mixed prior to transfer to the
autosampler vials.]
17.3.3 | Place black filter caps on top of the vials. |  | Use the tool provided to push the caps into the vials until
they are flush with the top |  | (see the IC manual for more detailed instructions).
17.3.4 | Wipe away any excess fluid from the top of the vial to avoid contamination from other samples.
17.3.5 | After all of the trays are filled, place them into the left side of the autosampler. |  | The white dot on
the tray indicates the first sample. Press the button labeled RUN/HOLD to the RUN position. The trays should
move until the first sample is under the sampling head. |  |  | The front panel should indicate a READY message.
Press local/remove switch to remove.
17.4 | Basic System Operations - Start-up and Shut-down
[Note: This procedure is specific for one instrument. Other instruments can be used with these procedures.]
17.4.1 | Start-up Procedure for Ion Chromatograph.
| 17.4.1.1 | The major components of the Dionex 2020i Ion Chromatography system are illustrated in
Figure 14. Turn helium and nitrogen tanks on by opening the valve on top of each tank (pressure in either tank
should not be less than 500 psi.). |  | Replace if necessary. | Open valves at the outlet end of both regulators.
Adjust pressure on the nitrogen regulator to 100 psi. |  |  | Adjust pressure on the helium regulator to 14 psi.
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-21

## Page 26
Name: “Table 1 of 1 on page 26 of 68”
Table: 26
Method IO-4.2 |  |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  |  | Atmospheric Acidic
17.4.1.2 Check the level of eluents and regenerating solutions. Turn the chromatography (CMA) values
for the anion channel switch ON. |  |  | Verify that the pressure reading on the face of the degassing unit is 7 psi.
Adjust by turning dial next to pressure gauge. |  |  |  | Turn the degas switch to HIGH.
| 17.4.1.3 |  | Turn the eluent reservoir switches, corresponding to the eluents to be degassed, to the ON
position. |  | Let the eluents degas on HIGH for 3-5 min, then turn degas switch to LOW.
| 17.4.1.4 |  | Select the appropriate program on the gradient pump module using the PROGRAM switch.
(Programs are recalled from memory by first pressing the PROGRAM switch, then the single digit reference
number corresponding to the appropriate program.)
| 17.4.1.5 |  | Prime the eluent lines.
[Note: | All of the eluent lines used during analysis must be primed to remove any air bubbles that may be
present. The selected program identifies which lines are used.]
| • | Open the gradient pump drawer. Turn the pump to the START position for 10 s, or until a CLICK
is heard. Turn the pump OFF. This step opens the valve to the eluent line displayed on the front
panel.
| • | Attach a 10 mL syringe to the priming block on the face of the gradient pump module. With the
priming block valve closed, pull the syringe plunger out to the end of the syringe.
| • | Open the priming block valve. |  | The syringe will quickly fill with eluent. | Close the valve on the
priming block when the syringe is almost full. Remove syringe from block and discard collected
eluant.
This priming procedure can be repeated if necessary. All of the eluent lines that are to be used during a day
of analysis should be primed at this time.
| 17.4.1.6 |  | Open the door of the Advanced Chromatography Module. | On the back of the door, at the
bottom, is the conductivity detector. |  |  | Four labeled lines (anion, cation, waste, and cell) are located next to the
cell. | The plumbing must be configured according to the type of analysis to be performed. |  |  |  | If anions are being
analyzed, attach the ANION line to the CELL line, and the CATION line to the WASTE line. |  |  |  |  | If cations are
being analyzed, attach the CATION line to the CELL line, and the ANION line to the WASTE line. Attach the
line coming from the pump to the correct port on the advanced chromatography module. |  |  |  |  | SYSTEM 1 on the
left is for anions; SYSTEM 2 on the right is for cations.
[Note: | If switching from one system to the other, the pump and the lines coming from the pump must be
purged of the original eluent, which is done by disconnecting the pump line from the chromatograph module,
turning the pump on, and running the new eluent into a waste beaker for 2-3 min.]
| 17.4.1.7 |  | Select the columns to be used (labeled pH or NO2 ) by pressing the blue button located below
the labels. |  | To verify that the correct columns are being used, press the switch should at least once, and then
set to the appropriate position.
| 17.4.1.8 |  | Turn the power switch on the autosampler ON (switch is located on the back of the unit, on
the right). |  | The default settings will be displayed on the front panel. |  | Attach the SAMPLE OUT line from the
autosampler to the advanced chromatography module. |  |  |  | The connection should be made to the port marked
SAMPLE of the appropriate system. |  |  | Turn the pump to START.
| 17.4.1.9 |  | Turn the conductivity cell ON (switch is located on the gradient pump module). |  | Turn the
REGEN switch for the appropriate system ON. |  |  |  | Verify that regenerant is flowing by inspecting the regenerant
waste line which empties into the sink. Open the advanced chromatography module door and inspect for leaks
at columns, fittings, etc. |  |  | Shut pump off if leaks are found.
Page 4.2-22 |  |  | Compendium of Methods for Inorganic Air Pollutants |  | June 1999

## Page 27
Name: “Table 1 of 1 on page 27 of 68”
Table: 27
Chapter IO-4 |  |  |  |  | Method IO-4.2
Atmospheric Acidic |  |  |  | Acidic/Basic Constituents
| 17.4.1.10 | Turn stripchart recorder ON. | Baseline should stabilize in less than 20 min. If baseline is not
stable, see troubleshooting Section 17.5 for assistance.
17.4.2 | Data acquisition start-up. | The following is a description of the current data acquisition program
used by the EPA. |  | The program is available (EPA, Atmospheric Chemistry and Physics Division, Office of
Research and Development, Research Triangle Park, NC) and is for IBM or IBM compatible computers. Other
appropriately designed programs may be used to compile the data collected for any given sampling network.
A computer programmed integrator does not need to be used to compute the data, but is recommended for large
sampling networks.
17.4.2.1 Turn on the IBM XT computer. From the C:>prompt, type: cd/cchart, then type: cchart. This
loads the Chromatochart software. |  | Turn switch on relay box to ENABLE, indicator light could go on.
| 17.4.2.2 Press F2 to enter the methods development module. |  | Select option number 1 - "select channel
# and load method file." |  | "Select channel # " type 0 or press ENTER to select the default choice shown in
the brackets (in this case 0). |  | "Load method file named" type the name of the appropriate method, then press
ENTER. | A directory of all of the current methods in memory can be obtained by pressing the F2 function key.
| 17.4.2.3 | Press F3 to enter the Data Acquisition module. | At this point you will be asked to save the
method file. | If there has not been any changes to the methods file, it does not need to be saved. Select option
#4 - "Collect Data." Press ENTER to deactivate the method queue. "Load Run Queue named," type the name
of the run queue if one has been created. |  | Type ENTER to deactivate the run queue.
| 17.4.2.4 | "Total # runs for method ," type how many times the method is to be repeated (total
number of samples). |  | "Autoanalyze Data" type Y. | "Autosave data to | disc" type Y. | "Data file name (xxxxx)
change?", type data file name. |  | "Press ENTER to begin methods." | Press ENTER only after the samples have
been loaded into the autosampler and the baseline has stabilized.
17.4.2.5 Figure 15 illustrates the chromatograms for each of the samples as output by the programmed
Spectra-Physics integrator.
17.4.3 | Calibration of IC. | The instrument should be brought to normal conditions with a warm-up time
of at least 30 min.
17.4.3.1 With the "Reading" light on, check to ensure the flow rate is 1.5 mL/minute, the fluid pressure
is 600 psi ± 100 psi and the conductivity is constant as measured by offset difference.
17.4.3.2 Fill the IC vials with the prepared standard solutions and (20, 10, 5 and 1 ppm SO4= and NO 3 G)
and pure eluent. |  | This will allow a four-point calibration curve to be made.
[Note: For low-level applications, more standards and blanks may be necessary in order to obtain accurate
reference curves.]
| 17.4.3.3 | Load the four vials into the sample vial holder and place the holder in the automated sampler
tray.
| 17.4.3.4 The tray is controlled by a Spectra-Physics SP4200 or SP4270 Computer Integrator. |  |  |  | Use the
integrators operation manual to begin calibrating. By using the RUN command, the analysis and data treatment
phases of the calibration are set in motion. |  | Four calibration standards are run, the chromatograms and peak
areas displayed for each run, and the run results for each anion are fitted to a quadratic curve by a least squares
regression calculation. The three curves are plotted and the correlation coefficients are calculated. The values
of the coefficients are normally greater than 0.999, where 1.000 indicates a perfect fit. Values of less than 0.99
indicate the calibration procedure should be repeated.
[Note: Recalibration should be carried out whenever standard concentrations show consistently high or low
results relative to the calibration curve is compared to the calibration curve from the old standards.
Comparability of points should be within ± 0.1 ppm or ± 10%. For standard concentrations of greater than
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants |  |  | Page 4.2-23

## Page 28
Name: “Table 1 of 1 on page 28 of 68”
Table: 28
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
1 ppm, comparability will normally be within 5% or better. Old standards are assumed correct since they
are referenced to the entire historical series of previous standard solutions all of which are comparable.]
17.4.4 |  | System Shut-down.
| 17.4.4.1 | Shut off the pump. | Turn the REGEN switch and the conductivity cell to the OFF position.
| 17.4.4.2 | Switch the eluent degas switch to HIGH.
| 17.4.4.3 Turn the stripchart recorder OFF. |  | Cap the pen. | Press the F10 function key on the computer.
Select option 3 to exit to DOS. |  |  | Shut off the printer and the computer.
17.4.4.4 Shut the eluent degas system and reservoir switches and the autosampler to the OFF position.
Close the valves on both gas cylinders. |  |  | Close the regulator valves.
17.5 | Basic Troubleshooting
Before proceeding with the troubleshooting guide, make sure the reagents used were prepared correctly and
are not "old."
17.5.1 |  | Unstable Baselines
| 17.5.1.1 Wavy baseline. |  | The most common reason for a wavy baseline is an air bubble in the gradient
pump. | This problem is diagnosed by observing the pump head indicator lights on the gradient pump module
front panel. |  | If the baseline is pulsing in phases with pump pistons, a bubble is usually present. |  | Other
possibilities include a dirty or stuck check valve, piston seal or "O"-ring, as well as an air bubble in the
conductivity cell.
17.5.1.2 Drifting baseline. Steadily increasing or decreasing baselines usually indicate that the suppressor
column is not performing as it should. |  |  | Parameters to change include the regenerant and eluent concentrations
and flow rates. |  | Check temperature routinely, as changes in temperature can cause drifting. |  | Balancing these
should stabilize the baseline if the suppressor is functioning correctly. |  |  |  | The Dionex manual describes clean-up
procedures if the suppressor is believed to be contaminated.
| 17.5.1.3 | High baselines. | As with drifting baselines, the parameters to change are eluent and regenerant
concentrations and flow rates. |  |  | A high baseline usually indicates that there is not enough baseline suppression;
this condition can be controlled by increasing the regenerant flow rate.
| 17.5.1.4 | Low baselines. | Low baselines usually indicate that there is too much suppression, which can
be controlled by decreasing the flow of the regenerant.
17.5.2 |  | Backpressure. | Variations in system backpressure are common and should not raise concern
UNLESS the pressure change is greater than 200 psi.
| 17.5.2.1 | High backpressure. | The system is protected from pressure related damage through the high
and low pressure alarm settings on the front panel of the gradient pump module. |  |  |  | If the high pressure setting
is correctly selected (200 psi above normal operating range), the pump will automatically shut-off if this value
is exceeded. |  | The reason for high backpressure is blockage in the system. |  | Possibilities include loading against
a closed valve, a plugged line, contaminated columns, etc. |  |  |  | Diagnosis is done by removing one component of
the system and observing how the pressure changes.
| 17.5.2.2 | Low pressure. | Low pressure readings usually indicate a leak somewhere in the system.
Carefully check all fittings for leaks. |  |  | Tighten if necessary.
17.5.3 |  | Flow
| 17.5.3.1 | Regenerant lines. | If there is no flow at the waste outlet end of the regenerant line, check the
following:
| • | Make sure the correct regenerant switch is turned on
| • | Verify that the reservoir is not empty
| • | Make sure the nitrogen tank is turned on
| • | Check that the regulator is correctly set
Page 4.2-24 |  |  | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 29
Name: “Table 1 of 1 on page 29 of 68”
Table: 29
Chapter IO-4 | [7 empty] | Method IO-4.2
Atmospheric Acidic | [5 empty] | Acidic/Basic Constituents
| 17.5.3.2 | Eluent lines. | If there is no flow at the outlet end of the eluent lines, check the following:
| • | Check that the pump is on
| • | Check that the eluent lines are connected to the correct port
17.5.4 |  | Software. | Refer to the ChromatoChart manual for detailed information on software problems.
18. Ammonia Analysis by Technicon Autoanalysis
Presented in Sections 18.1 and 18.2 are the recipes for the standards and reagents required to analyze the
ammonium ion (NH4+ or ammonia (NH 3)) by Technicon autoanalysis. The prelude of these sections briefly
describes the TRAACS 800 autoanalyzer and the sample flow through the TRAACS 800 for NH 4+ analysis.
This instrument is capable of quantifying, from a single sample, three different species simultaneously. An
aliquot of the sample is taken from an automated sampler by syringe. A splitter divides the aliquot into the
appropriate volumes required for the particular analyses. |  |  | Each of the volumes is then transferred to the
appropriate analytical cartridge. |  |  | Sample flow diagrams that illustrate SO 4= , NO3- and NH4+ analysis
(although in Method IO-4, only ammonia is analyzed by this technique) from the citric acid denuder and
Teflon® filter can be shown separately and independently of one another. | [5 empty] | The data computation (by
computer) and quality assurance protocols, however, can not be readily adapted to single-channel
instruments. |  | These protocols need to be specific to the individual analytical instrument. | [5 empty] | In brief, NH 4+
analysis is illustrated in Figure 16. The samples, along with all standards, are taken from the auto-advance
sampler tray by the use of a proportioning pump and automated syringe. Air and EDTA are first added to
the samples and are mixed in the first set of coils. After mixing, phenolate is added and mixed in the next
set of coils. Nitroprusside is then added and mixed, followed by the addition and mixing of hypochlorite.
At this stage, the sample should be a bright blue color. After the last mixing stage, the sample is sent through
a heated bath, followed by another mixing stage. Finally, the sample is sent through a colorimeter where the
results are recorded on a digital printer and stored in a computer file for further manipulation.
[Note: Ammonia can be analyzed by other techniques.]
18.1 | Standards and Stock Solutions Preparation
[Note: | Before discarding the old solution, it should be checked against the fresh solution by comparing
calibration curves on the working solutions prepared from them. |  |  |  | Slopes and intercepts are calculated for
each set of standards. The old slope and intercept are used to calculate concentration values from readings
for the new standards. This calculation determines if the old solution has deteriorated or if an error has been
made in preparing the new solution.]
18.1.1 |  | Ammonium Solution Standard (1,000 μg/mL). | Dry ammonium chloride in an oven for 1 h at
50-60EC and desiccate over silica gel for 1 h. |  |  | Weigh 2.9470 g ammonium chloride and dissolve in 800 mL
DDW. | Dilute to 1 L with DDW and mix thoroughly. |  | This solution is stable for 1 y.
18.1.2 |  | Intermediate Ammonium Standards. | To make a 100 μg/mL ammonium standard, pipet 10 mL
of ammonium stock standard into a 100 mL volumetric flask. Dilute to volume with DDW and mix thoroughly.
Keep refrigerated. This solution remains stable for 1 month. To make a 10 μg/mL ammonium standard, pipet
1.0 mL of ammonium stock standard into a 100 mL volumetric flask. |  |  |  | Dilute to volume with DDW and mix
thoroughly. |  | This solution remains stable for one week.
18.1.3 |  | Working Ammonium Standards in DDW. | Pipet | aliquots | of | the | 100 | μg/mL | ammonium
intermediate standards into 100 mL volumetric flasks according to the table below. | [6 empty] | Dilute to volume with
DDW. | Prepare fresh daily.
June 1999 |  |  | Compendium of Methods for Inorganic Air Pollutants |  |  |  |  | Page 4.2-25

## Page 30
Name: “Table 1 of 1 on page 30 of 68”
Table: 30
Method IO-4.2 |  |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
|  | Stock or intermediate |  | Concentration,
| Standard | standard, Fg/mL | Aliquot, mL | Fg/mL
| A | 1,000 | 4.0 | 40.0
| B | 100 | 4.0 | 4.0
| C | 100 | 3.0 | 3.0
| D | 100 | 2.0 | 2.0
| E | 100 | 1.0 | 1.0
| F | 100 | 0.5 | 0.5
| G | 10 | 2.0 | 0.2
| H | 10 | 1.0 | 0.1
18.2 | Reagent Preparation
[Note: When reagents are prepared, label the container with the contents, concentration, date prepared, and
the preparer's initials.]
18.2.1 | Alkaline Phenol. | Add 83.0 g loose crystallized phenol to 800 mL DDW in a 1 L volumetric flask.
Keeping the flask in an ice bath or under tap water, slowly add 96.0 mL 50% sodium hydroxide solution. Shake
the flask while adding the sodium hydroxide. |  | Cool to room temperature, dilute to 1 L with DDW, and mix
thoroughly. | Store in an amber glass container. | This solution will remain stable for 3 mo, if kept out of direct
light.
18.2.2 | Sodium Hypochlorite Solution. | The amount of sodium hypochlorite solution varies from batch to
batch of sodium hypochlorite (5% commercial grade). |  |  | Therefore, for each new batch, a base and gain
experiment must be run to adjust the amount of sodium hypochlorite required to obtain the existing base and
gain values. | In a 150 mL volumetric flask, dilute 86 mL of 5% sodium hypochlorite solution to 100 mL with
DDW and mix thoroughly. Check base and gain values. Reduce or increase the amount of sodium hypochlorite
to obtain the same base and gain values as the previous sodium hypochlorite batch. This solution remains stable
for 1 day.
18.2.3 | Sodium Nitroprusside Solution. | Dissolve 1.1 g of sodium nitroprusside in about 600 mL of DDW,
dilute to 1 L with DDW, and mix thoroughly. |  | Store in an amber container and keep in refrigerator. |  |  | This
solution remains stable for 1 month, if kept out of direct light.
18.2.4 Disodium EDTA Solution. Dissolve 1.0 mL of 50% w/w sodium hydroxide and 41.0 g of disodium
EDTA mix thoroughly. | Add 3.0 mL of Brij-35 and mix. |  | Store in plastic container. | This solution will remain
stable for 6 months.
19. pH Analysis
19.1 | Standard and Reagent Preparation
[Note: Each of the standard H 2SO 4 stock solutions must be prepared fresh the day of pH analysis.]
Page 4.2-26 | Compendium of Methods for Inorganic Air Pollutants |  |  |  | June 1999

## Page 31
Name: “Table 1 of 1 on page 31 of 68”
Table: 31
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
19.1.1 | Standard H2 SO4 | Solution.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-27

## Page 32
Name: “Table 1 of 1 on page 32 of 68”
Table: 32
Method IO-4.2 | [5 empty] | Chapter IO-4
Acidic/Basic Constituents | [5 empty] | Atmospheric Acidic
| 19.1.1.1 | Label seven 25 mL polyethylene stoppered volumetric flasks. |  |  | Also, label each flask with the
volume of 1 N H2 SO4 |  | solution indicated in the following table:
Volume of
|  |  | 1.000 N H2SO 4 |  | Standard
|  | Standard H 2SO 4, | added to each |  | concentration,
|  | Flask No. | flask, μL |  | 10 -3N H2SO 4
|  | 1 |  | 0 | 0
|  | 2 |  | 25 | 1
|  | 3 |  | 50 | 2
|  | 4 |  | 100 | 4
|  | 5 |  | 200 | 8
|  | 6 |  | 400 | 16
|  | 7 |  | 800 | 32
| 19.1.1.2 | Use the 25 μL automatic pipet to add 1 N stock H2 SO4 |  |  | to flasks #1-3. | Use the 100 μL pipet
to add 1 N stock H 2 SO4 |  | to flasks #4-7. | Dilute all flasks to the 25 mL mark with absolute ethanol. |  |  | Cap with
stoppers or parafilm and mix well.
19.1.2 | 2 M Potassium Chloride (KCl) Solution.
| 19.1.2.1 | Weigh 149.2 ± 0.1 g of KCl. | Add the KCl to a 2 L flask.
| 19.1.2.2 | Add about 700 mL of DDW water to the flask. |  |  | Swirl the solution until the KCl is completely
dissolved.
| 19.1.2.3 | Pour this mixture into a 1 L graduated cylinder. |  |  | Rinse the flask with a small amount of water
and transfer the rinse into the cylinder. |  |  | Fill the cylinder to the 1 L mark.
| 19.1.2.4 Pour the solution from the cylinder into the 1 L polyethylene bottle. |  |  |  |  | Cap and shake the bottle
to mix well. | Mark the bottle with date of preparation.
19.1.3 | 0.1 N Perchloric Acid (HClO4 ) Solution.
19.1.3.1 Fill a 1 L graduated cylinder about 1⁄2 full with DDW. Transfer 10 ± 0.1 mL of 60-62% HC104
into the 1 L cylinder with a 10 mL pipet.
| 19.1.3.2 | Fill the cylinder to the 1 L mark. | Pour the solution into the 1 L polyethylene bottle.
| 19.1.3.3 | Cap and shake the bottle to mix well. |  | Mark the date of preparation on the bottle.
19.1.4 | 0.01 N HClO 4 | Solution.
| 19.1.4.1 | Fill a 1 L graduated cylinder about 1⁄2 full with DDW.
| 19.1.4.2 | Measure 100 mL of the 0.1 N HC1O 4 |  |  | solution with the 100 mL graduated cylinder. | Add
solution to the 1 L cylinder.
| 19.1.4.3 | Fill a 1 L cylinder with DDW to the 1 L mark. |  |  | Pour the solution into the 1 L polyethylene
bottle.
| 19.1.4.4 | Cap and shake the bottle to mix well. |  | Mark the date of preparation on the bottle.
19.1.5 | Extraction Solution (0.0001 N HClO 4 ).
[Note: This solution has the same composition as the solution used to fill the sample vials for Teflon® filters.
It must be prepared fresh on the day of pH analysis.]
| 19.1.5.1 | Using a 5 mL calibrated automatic pipet, add 10 ± 0.1 mL of 0.01 N perchloric acid (HClO 4 )
to a 2L Erlenmeyer flask of water. |  |  | Add 990 ± 10 | mL of DDW to the flask.
Page 4.2-28 |  | Compendium of Methods for Inorganic Air Pollutants |  |  |  | June 1999

## Page 33
Name: “Table 1 of 1 on page 33 of 68”
Table: 33
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
| 19.1.5.2 | Mix well and cover with parafilm until ready for use.
19.1.6 | EA Solution.
19.1.6.1 Measure 150 ± 2 mL of 0.0001 N HClO 4 (prepared in Section 19.1.5) into a 250 mL graduated
cylinder. | Transfer to a 250 mL Erlenmeyer flask.
| 19.1.6.2 | Using a 5 mL graduated cylinder, add 5 ± 0.1 mL of ethanol (from the same fresh bottle of
ethanol that was used to prepare the standards in 19.1.1) to the flask.
| 19.1.6.3 | Again using a 5 mL graduated cylinder, add 3 ± 0.1 mL of 2 M potassium chloride (KCl)
solution (see 19.1.2) to the flask.
| 19.1.6.4 | Mix well and cover with parafilm until ready for use.
19.1.7 | Working Standard Test Solutions.
| 19.1.7.1 | Place fourteen-4 mL polystyrene sample cups (as used with Technicon Auto-Analyzer II
system) labeled 1, 1*, 2, 2*...7, 7* into racks. |  | Using the calibrated dispensing pipet bottle, add 3 mL of EA
solution to each 4 mL cup.
| 19.1.7.2 | Using the displacement pipet, add 50 μL of absolute ethanol to each cup. | Pour about 3 mL
of standard (H 2 SO4 |  | solution) #1 into a labeled 4 mL cup.
| 19.1.7.3 | Immediately pipet 50 μL of this standard into the 4 mL cups labeled 1 and 1* containing the
EA solution and ethanol.
[Note: | This transfer must be done without delay to prevent the standard concentration from increasing
significantly due to evaporation of the ethanol solvent.]
| 19.1.7.4 | Repeat the procedure for each of the other 6 standards. If there is a delay of more than 5 min
between the preparation of these mixtures and the next step, put caps on the 4 mL cups.
| 19.1.7.5 To prepare for analysis, mix each mixture. |  | Then transfer two aliquots from each cup to 2 mL
sample cups. | Place cup #1 in a rack. | In a second rack place two-2 mL cups labeled 1 and 1-. | Use the 1 mL
automatic pipet to mix the contents of 4 mL cup #1 by drawing 1 mL into the pipet tip and then dispensing it
back into the 4 mL cup three times. |  | Use the same pipet to transfer 1 mL of the mixture to each of the two
labeled 2 mL cups. |  | Place caps on the two 2 mL cups. | After transferring the two aliquots to 2 mL cups, rinse
the automatic pipet tip in a flask of DDW. Repeat the transfer procedure for each of the other working standard
pairs.
19.2 | Calibration of pH Meter
The pH meter requires temperature calibration whenever a new electrode is used. |  |  | Use the manufacture's
procedure in the instrument manual. This calibration should be repeated every 3 months when not in use.
The pH meter is left with the power cord plugged into the AC outlet, the mode control knob is left in the
standby position, the electrode lead is partially disconnected by pressing the plastic ring on its outer edge,
and the combination electrode is immersed in a 4 M KCl solution (a slit rubber stopper seals the bottle with
the electrode in it). Keep a record of the temperature calibrations in a lab notebook.
19.3 | Pre-Analysis Calibration
[Note: The pH buffer solutions are not used for any quantitative purpose. They are used to standardize the
electrode and as a diagnostic to verify that the pH measurement system is working as expected before
beginning analysis of the samples.]
19.3.1 | Use a pH Analytical Log Form to record all data (see Figure 17). |  | While still in standby mode,
reconnect the electrode lead at the back of the pH meter.
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-29

## Page 34
Name: “Table 1 of 1 on page 34 of 68”
Table: 34
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
19.3.2 | Fill three 4 mL cups with pH 7 buffer. |  | Withdraw the electrode from the 4 M KCl bottle and wipe
the tip gently with a Kimwipe® to remove the bulk of the solution. |  |  | Rinse the electrode with one cup of pH 7
buffer. | Do not test pH of the first cup.
19.3.3 | Immerse the electrode in the second cup of the pH 7 buffer. |  |  | Use a small bottle or other support to
hold the cup up to the electrode while waiting for the meter reading to equilibrate.
19.3.4 Test the pH by turning to the pH mode of the meter. |  |  | Allow the reading to stabilize for at least 30 s.
Record the result on the log for "1st cup."
19.3.5 | Turn to standby mode, and then test the last cup of pH 7 buffer. |  |  | Record the results on the log for
the "2nd cup." | If the pH value for the 2nd cup is not 7.00 ± 0.01, adjust the "calib." knob to obtain a reading
of 7.00. | Note this adjustment on the log.
19.3.6 | Fill three 4 mL cups with pH 4 buffer. |  | With the meter in the standby mode, remove the cup
containing pH 7 buffer, wipe the tip of the electrode gently with a Kimwipe® and then rinse the electrode with
the first cup of pH 4 buffer.
19.3.7 | Test the next two cups of pH 4 buffer as above, recording the results on the log. |  |  | If the pH value
for the 2nd cup is not 4.00 ± 0.01, adjust the "slope" knob to get a reading of 4.00. |  |  |  | If the value for the second
cup was not 4.00 ± 0.03, the calibrations at pH 7 and at pH 4 must both be repeated.
19.4 | pH Test 0.01 N HClO4 | Solution
[Note: | The 0.01 N HClO 4 solution is used to prepare the ES solution, which is used to prepare the EA
solution. The pH value for the EA solution must be 4.09 ± 0.04. If this pH value is not achieved, the 0.01
N HClO 4 solution must be reprepared.]
19.4.1 | Calibrate the pH meter with pH 4 buffer.
19.4.2 | Rinse the pH electrode with DDW. |  | Wipe the tip of the electrode with a Kimwipe®.
19.4.3 | Fill three 4 mL cups with EA solution. |  | Measure the pH of the test EA solution as with the buffer
solutions this value must be 4.09 ± 0.10.
19.4.4 | If the above pH value is not achieved, follow the steps outlined in Sections 18.1.3 through 18.1.6
to reprepare the solutions. | Test the pH of the new solutions. |  | Repeat as necessary to obtain a pH of 4.09 ±
0.04.
19.4.5 | Leave the electrode immersed in the "2nd cup" with the meter in the standby mode until ready to
start analysis of the working standards.
19.5 | Analysis of Working Standard
[Note: Immediately following the EA analysis, start testing the working standards.]
19.5.1 With the pH meter still in the standby mode, remove the last cup from the electrode, gently wipe the
tip with a Kimwipe®, and immerse the electrode into the working standard cup #1.
[Note: | Only two cups are available for each working standard (also for filter extracts). |  |  | Thus, pH
measurement is made for both of the two cups for each sample. Also, the electrode tip is not wiped between
the 1st and 2nd cups of each sample.]
19.5.2 | After testing the pH of cup #1, test cup #1-. |  | Record the results of both on the pH Analytical Log
Form.
19.5.3 | With the meter in the stand-by mode, remove the #1- sample cup, wipe the electrode with a
Kimwipe®, and test one 2 mL cup of EA solution. |  |  | Rinse with DDW.
Page 4.2-30 | Compendium of Methods for Inorganic Air Pollutants |  |  | June 1999

## Page 35
Name: “Table 1 of 1 on page 35 of 68”
Table: 35
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
19.5.4 Test a 2nd cup of EA solution; record the results for both cups on the logsheet. |  |  | Discard the 1st cup
of EA, but retain the 2nd cup to be used as the 1st cup for the next EA test.
19.5.5 Continue testing the remainder of the working standards, #1*, 1*-, ... 7, 7-, 7*, 7*-. Remember that
the electrode tip is wiped before and after each pair of test solutions, but not in between two cups of the same
sample.
[Note: | If there is trouble obtaining constant pH values, use a magnetic stirrer to keep the contents to be
measured uniform. If employed, ensure that the sample cups are insulated from any temperature increase of
the stirring platform, which may occur during extended use.]
19.5.6 |  | Use the mode control knob in the "temp." position to measure the temperature of the test solutions
every 5-10 samples and record the results on the logsheet.
19.6 | Analysis of Filter Extracts
After measuring the pH of the working standards, measure the pH of the filter extracts to which 20 FL of 2M
KCl solution (19.1.2) has been added, and record all results on the log. After all the filter extracts have been
tested make an additional test with the EA solution. At the end make a final test of pH 4 buffer. With the
mode control in the standby mode, shut down the pH meter by disconnecting the electrode lead at the back
of the meter, leaving the meter power cord plugged into the AC line. Immerse the electrode tip in the bottle
of 4 M KCl.
20. Atmospheric Species Concentration Calculations
The system described in the previous sections collects nitric acid (HNO 3), sulfur dioxide (SO 2), ammonia
(NH3), particulate sulfate (SO 4= ), particulate nitrate (NO 3- ), particulate ammonium (NH4+ ), and acidic (H + )
particles. The collection of each of these species is illustrated in Figure 1. |  |  | Nitric acid and sulfur dioxide
gases are collected on denuder one. Ammonia gas is collected on denuder two. Particulate sulfate, nitrate,
and ammonium are collected on the first (Teflon)® filter, while some of the particulate nitrate collected on
the Teflon® filter evaporates and collects on the second (nylon) filter. Also collected on the Teflon® filter
are fine particles that contain hydrogen ions (H+ ), though probably not free H + . Hydrogen ions are most
likely present in the H 3O + form. The concentration of these H+ ions indicates the atmospheres acid aerosol
content and is collected on the nylon filter. |  | Prepare the Teflon® filter extracts for pH analysis prior to IC
analysis for the particulate sulfate contents. Special precautions must be taken to prevent contamination of
the Teflon® filters by ammonia before either of the analyses.
20.1 | Assumptions of the Annular Denuder System
20.1.1 |  | A number of assumptions are made about the performance of the annular denuder system so that
validity of the calculations presented later in this section will hold true. |  | As discussed in Section 6, significant
interferences need to be considered to make accurate estimations of species concentrations. |  |  | The assumptions
are as follows:
| • | The first denuder stage collects 100% of sampled HNO 3 as nitrate. (Since the diffusivity of HNO3
is high, diffusion to the side walls is assumed to be very quick.)
| • | The first denuder collects 100% of the SO 2 as sulfite, which can oxidize to sulfate.
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-31

## Page 36
Name: “Table 1 of 1 on page 36 of 68”
Table: 36
Method IO-4.2 | [5 empty] | Chapter IO-4
Acidic/Basic Constituents | [5 empty] | Atmospheric Acidic
[Note: Before analysis, add hydrogen peroxide (H 2O 2) to oxidize the sulfite (SO 3- ) to sulfate (SO 4= ) to
simplify the calculations.]
| • | The second denuder stage collects 100% of the sampled ammonia (NH 3) as ammonium ion (NH 4+ ).
| • | The Teflon® filter is 100% efficient for particulate sulfate, nitrate, and ammonia. Particle losses
are less than 1% on each denuder. This assumption may or may not stand true depending on the
|  | concentrations of the components in the air sampled. |  |  |  | Modifications may be needed to avoid low
|  | (or underestimates of) acidic measurements. |  |  | For example, another filter stage may need to be
added to accurately account for the particulate ammonia content of the air sampled. If ammonium
nitrate (NH4NO3) was collected on the Teflon® filter, its probability of evaporation is high.
Therefore, a citric acid-impregnated filter downstream would correct for the loss from the Teflon®
filter. Also, interaction of ammonia and sulfuric acid neutralizes the filter and causes the acidic
measurement to be biased. (Again diffusion rules the particle loss assumption; particles have lower
diffusivities than gases.)
| • | The nylon filter collects any nitrate that evaporates from the Teflon® filter.
20.1.2 |  | Analytical results from the IC and technician autoanalysis are given in units of μg/mL from NO3- ,
SO4=, and NH 4+, respectively.
20.1.3 |  | The extraction volume was 11 mL for Na2 CO3 |  |  | denuder(s) and 10 mL (i.e., 0.010 L) for the citric
acid denuder extractions and the nylon filter and 6.2 mL for the Teflon® filter.
20.2 | Calculation of Air Volume Sampled, Corrected to Standard Conditions
20.2.1 The total sample air volume, Vt , for each sample is calculated using the data from the Field Test Data
Sheet. | These data include the initial and final elapsed times, the initial rotameter reading, and the rotameter I.D.
No. | Use the calibration curve for the given rotameter to calculate the flow for the sample, in LPM, if
applicable. |  | Calculate the value of Vt |  | as follows:
V t = [F] [t]
where:
F = flow from the calibration curve, L per minute.
t = net elapsed time, min.
V t = total sample volume, L.
20.2.2 |  | Convert L to m3 | by:
V s = V t x (10 -3)
where:
V s = total sampling volume, m 3.
10 -3 = conversion factor, m 3/L.
20.2.3 |  | Calculate the air volume sampled, corrected to EPA-reference conditions:
Page 4.2-32 |  | Compendium of Methods for Inorganic Air Pollutants |  |  |  | June 1999

## Page 37
Name: “Table 1 of 1 on page 37 of 68”
Table: 37
Chapter IO-4 | [9 empty] | Method IO-4.2
Atmospheric Acidic | [9 empty] | Acidic/Basic Constituents
[6 empty] | T | P
|  |  | V |  | V | Y(std | )(bar | )
|  |  | std | ' | s
[6 empty] | T | P
[6 empty] | m | std
where:
| V | volume of sample at EPA-reference conditions, m3 .
std
| Vs = | volume of gas sample through the dry gas meter, or calculated volume sampled as
indicated by rotameter (see Section 20.2.2), m3 .
| T std = | absolute EPA-reference temperature, 298EK.
| T m = | average flowmeter or dry gas meter temperature, EK.
| P bar | barometric pressure of flow or volume measurement condition, mm Hg.
| P std = | EPA-reference barometric pressure, 760 mm Hg.
| Y = | dry gas meter calibration factor (if applicable), dimensionless.
20.3 | Calculations of Concentration Using Results from IC and Technicon Autoanalysis
20.3.1 | Analytical Results for HNO 3 |  | from Na 2 CO 3 |  |  | Denuder
Cg (HNO 3 ), μg/m3
This factor 1.016 represents the ratio of molecular weights of HNO3- and NO 3- . Subscript "g" denotes "gas."
20.3.2 | Analytical Results for SO 2 |  | from Na 2 CO 3 |  |  | Denuder
Cg (SO 2 ), μg/m3
The factor 0.667 represents the ratio of molecular weights of SO2 | [8 empty] | and SO 4=. | Subscript "g" denotes "gas."
20.3.3 | Analytical Results for NH 3 |  | from Citric Acid Denuder
Cg (NH 3 ), μg/m3
This factor 0.944 represents the ratio of molecular weights of ammonia to ammonium ion. | [9 empty] | Subscript "g"
denotes "gas."
20.3.4 | Analytical Results for Particulate SO4= Captured on Teflon® Filter
|  | Cp | (SO 4 ), μg/m3 |  | [SO4= (μg/mL) x 6.2 mL]/Vstd
This formula expresses the assumption that essentially all of the particulate sulfate is collected on the
Teflon® filter and no evaporation occurs. |  |  | The subscript "p" denotes "particles."
20.3.5 | Analytical Results for Particulate NH4+ Captured on Teflon® Filter
|  | Cp | (NH 4+), μg/m3 |  | [NH4+ (μg/mL) x 6.2 mL]/Vstd
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | [7 empty] | Page 4.2-33

## Page 38
Name: “Table 1 of 1 on page 38 of 68”
Table: 38
Method IO-4.2 | [10 empty] | Chapter IO-4
Acidic/Basic Constituents | [9 empty] | Atmospheric Acidic
This formula expresses the assumption that essentially all of the particulate ammonia is collected on this
Teflon® filter and no evaporation occurs. |  |  |  |  | The subscript "p" denotes "particles."
20.3.6 | Analytical Results for Particulate NO3G Captured on Teflon® Filter
|  | Cpt | (NO 3G p ), μg/m3 |  |  | [NO3 G (μg/mL) x 6.2 mL]/Vstd
Nitrates evaporator from the Teflon® filter during sampling, so one must also calculate the NO3G on the
nylon filter. | The subscript "pt" denotes particles on Teflon® filters.
|  | Cpn | (NO 3 G), μg/m3 |  |  | [NO3 G (μg/mL) x 10 mL]/Vstd
The subscript "pn" denotes particles on the nylon filter. | [6 empty] | Calculate the particulate nitrate concentration by:
|  |  | Cp | (NO 3 G) = Cpt |  | (NO 3 G) + Cpn |  | (NO 3 G)
20.4 | Calculations Using Results from pH Analysis
Earlier pH determinations have been based on the pH buffer concentrations, the activity of the solution, and
the antilog of the measured pH value. |  |  |  | More recent studies have steered away from the issue of activity by
comparing the results of the standards, thus alleviating errors introduced by basing the activities of ions retained
on filters on those retained in solution. The methodology developed from these more recent studies is described
herein. | The end results are reported in terms of mass of equivalent of ions. | [7 empty] | Appropriate values of accuracy
and precision with respect to H+ concentration for this method are 10% and 5%, respectively, for sample pH
values in the 4.00 to 7.00 range.
20.4.1 | Adjustment | for | Filter | vs. | Non-Filter | Standards. |  | This | adjustment | is | necessary | because
experiments showed that the measured acid concentration from filters doped with H 2 SO4 stock standards yielded
concentrations, as measured by the difference from EA solution, that were about 3% lower than the values
found for working standards (prepared without filters from the same stock standards). | [9 empty] | The results gave the
following relation (by linear regression):
Cf
where:
Cf
Cnf | [6 empty] | based on standards prepared without filters.
For each working standard (non-filter), on a given analysis day, calculate the "apparent net strong acid
concentration of H2 SO4 " as follows:
[6 empty] | -pHWS | -pHEA
[5 empty] | Cnf | - 10
Page 4.2-34 |  | Compendium of Methods for Inorganic Air Pollutants | [8 empty] | June 1999

## Page 39
Name: “Table 1 of 1 on page 39 of 68”
Table: 39
Chapter IO-4 | [11 empty] | Method IO-4.2
Atmospheric Acidic | [11 empty] | Acidic/Basic Constituents
where:
pHWS = measured pH for a working standard (or apparent strong acid concentration for non-filter H 2 SO4
doped standards).
pHEA = measured pH for the EA solution (or apparent strong acid concentration for non-filter, non-H2 SO4
doped standard).
After calculating the Cnf |  |  | values for each working standard, use the above equation to calculate the adjusted
values of Cf |  | for each.
| 20.4.2 | Determination of Nominal Strong Acid Concentration for Filter Samples. | [9 empty] | The apparent net
strong acid concentration of each sample filter extract, Cs, is calculated as with the working standards:
[8 empty] | -pHS | -pHEA
[7 empty] | Cs = 10 | - 10
where:
pHS = measured pH of the sample filter extract (or apparent strong acid concentration for sample filters
extracts).
| pHEA = measured |  | pH | for | the | EA | solution | (or | apparent | strong | acid | concentration | for | non-filter,
|  | non-H 2 SO4 | standards).
[Note: The Cs values for the filter extracts are directly comparable to the C f values for the working standards,
since the Cf values have been adjusted for the difference in apparent acid concentration for tests made with
filters and tests made without filters.]
| 20.4.3 Calculation of Strong Acidity Aerosol Concentration. | [8 empty] | Calculate the final concentration of
apparent net fine particle (
[7 empty] | C(H +) = Cf |  | / Vstd
where:
CH + = apparent net fine particle str\\\\\=ong acidity concentration, Fg/m3 .
Cf
Vstd = volume of sampled gas at EPA-referenced conditions (see Section 17.2.3), m3.
21. | Variations of Annular Denuder System Usage
As described in Sections 3 and 4, the ADS is used to measure reactive acidic (SO2 and HNO 3 ) and basic (NH3 )
gases, particulate sulfates and nitrates, and strong acidity of atmospheric particles (
air. | The unique features of the ADS that separate it from established air monitoring methods are the ability of
sampling artifacts to be eliminated from the collected gases and particles and preservation of the samples for
subsequent analysis, which is accomplished by removing NH3 | [9 empty] | in the gas stream with a citric acid coated
denuder, thus reducing the probability of the particulate acid sulfates (SO4=) captured on the Teflon® filter from
being neutralized. The ADS configuration described in this methodology clearly illustrates these unique features.
The elutriator is designed to allow only particles with | [11 empty] | The impactor is
designed to reduce the possibility of coarse particle infiltration even further. | [10 empty] | And finally, the sequence of the
June 1999 |  |  | Compendium of Methods for Inorganic Air Pollutants | [9 empty] | Page 4.2-35

## Page 40
Name: “Table 1 of 1 on page 40 of 68”
Table: 40
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
denuders reduces interference of possible chemical reactions that could cause under-or over-estimations of
concentrations to be made. |  | Although this configuration is recommended for measuring these gases and
particulates, the user may wish to measure only one or two of the chemical species. The following discussion
will present possible variations of the ADS to accommodate such usages.
21.1 Today, the ADS is being used in intercomparison studies to assess NH3 concentration differences indoors
and outdoors. | The assembly used here consists of an elutriator-impactor assembly, an annular denuder, and
a filter pack assembly. |  | The elutriator-impactor assembly and the annular denuder are both smaller than those
described earlier. The filter pack is available in the smaller size, but an adaptor is also available to assemble the
smaller annular denuder to the larger filter pack assembly. |  | This system is referred to as the personal sampler
(see Figure 18). | It is designed for sampling while attached to the shirt of a worker. |  | The personal sampler can
be used to measure other chemical species in indoor air by simply changing the reactive surface (coating) of the
annular denuder or the types of filters used.
21.2 Another variation of ADS application is simultaneous use in parallel with a fine particle sampler. The fine
particle sampler assembly is very similar to the annular denuder assembly. |  | The main difference is that a
flow-straightener tube replaces the annular denuder. The flow-straightener is a shorter version, 1-1/4 to 4" long,
of the annular denuder and creates even air flow across the filters to collect particulate matter. |  |  | Again the
elutriator-impactor assembly and flow-straightener are available in smaller sizes with accommodating filter pack
assemblies. | In addition, the ADS carrying and shipping cases as well as the sampling box can be adjusted to
accommodate the ADS and fine particle sampler. |  | The assemblies as they would appear in the sampling box
ready for sampling are illustrated in Figure 19.
21.3 If one has interest in quantitations HNO2 utilizing the ADS, special sampling and analytical concerns must
be addressed. | As identified in Section 24, Citation 14, guidance is given for accurate quantitation of HNO2 |  | in
ambient air.
21.4 | The annular denuder methodology has been extended to other constituents, as indicated in Table 5.
22. | Method Safety
This procedure may involve hazardous materials, operations, and equipment. |  | This method does not purport
to address all of the safety problems associated with its use. |  | The user must establish appropriate safety and
health practices and determine the applicability of regulatory limitations prior to the implementation of this
procedure. | These practices should be part of the user's SOP manual.
23. | Performance Criteria and Quality Assurance (QA)
Required quality assurance measures and guidance concerning performance criteria that should be achieved
within each laboratory are summarized and provided in the following section.
Page 4.2-36 |  | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 41
Name: “Table 1 of 1 on page 41 of 68”
Table: 41
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
23.1 | Standard Operating Procedures (SOPs)
23.1.1 | SOPs should be generated by the users to describe and document the following activities in their
laboratory: (1) assembly, calibration, leak check, and operation of the specific sampling system and equipment
used; (2) preparation, storage, shipment, and handling of the sampler system; (3) purchase, certification, and
transport of standard reference materials; and (4) all aspects of data recording and processing, including lists of
computer hardware and software used.
23.1.2 | Specific instructions should be provided in the SOPs and should be readily available to and
understood by the personnel conducting the monitoring work.
23.2 | QA Program
The user should develop, implement, and maintain a quality assurance program to ensure that the sampling
system is operating properly and collecting accurate data. Established calibration, operation, and maintenance
procedures should be conducted regularly and should be part of the QA program. |  |  | Calibration procedures
provided in Sections 17 and 19, operation procedures in Sections 14 and 17, and maintenance procedures in
Section 17 of this method and the manufacturer's instruction manual should be followed and included in the QA
program. | Additional QA measures (e.g., trouble shooting) and further guidance in maintaining the sampling
system are provided by the manufacturer. |  | For detailed guidance in setting up a quality assurance program, the
user is referred to the Code of Federal Regulations (Section 24, Citation 12) and the U. S. EPA Handbook on
Quality Assurance (Section 24, Citation 13).
23.2.1 Field QA. |  | It is recommended that the flow rates of each denuder system be audited at least
quarterly.
23.2.2 Laboratory Quality Control (QC). |  | It is recommended that laboratory QC program include the
following as minimum requirements.
23.2.2.1 Include a reagent blank with each set of twenty (20) samples or less of each matrix (i.e., denuder
extract, filter extract, etc.).
23.2.2.2 Include a laboratory duplicate with each set of twenty (20) samples or less of each matrix.
23.2.2.3 Include a laboratory control sample (LCS), also known as a laboratory blank spike (LBS) with
each set of twenty (20) samples or less of each matrix.
24. References
1. | Waldman, J. M., Operations Manual for the Annular Denuder System Used in the U. S. EPA/RIVM
Atmospheric Acidity Study, UMPNJ - Robert Wood Johnson Medical School, Piscataway, NJ, August 28,
1987.
2. | American Chemical Society Subcommittee on Environmental Chemistry, "Guidelines for Data
Acquisition and Data Quality Evaluation in Environmental Chemistry," Analyt. Chem., Vol. 52:2242-2249,
1980.
3. | Sickles, II, J. E., Sampling and Analytical Methods Development for Dry Deposition Monitoring,
Research Triangle Institute Report No. RTI/2823/00-15F, Research Triangle Institute, Research Triangle
Park, NC, July 1987.
June 1999 |  | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-37

## Page 42
Name: “Table 1 of 1 on page 42 of 68”
Table: 42
Method IO-4.2 | [9 empty] | Chapter IO-4
Acidic/Basic Constituents | [7 empty] | Atmospheric Acidic
4. Forrest, J., and Neuman, L., "Sampling and Analysis of Atmospheric Sulfur Compounds for Isotopic
Ratio Studies," Atmos. Environ., Vol. 7:562-573, 1973.
5. | Stevens, R. K., et al., ACGIH Symposium: |  |  | "Inlets, Denuders and Filter Packs to Measure Acidic
Inorganic Pollutants in the Atmosphere," Aislomer Conference Center, Pacific Grove, CA, February 16,
1986.
6. Appel, B. R., Povard, V., and Kothney, E. L., "Loss of nitric acid within inlet devices for Atmospheric
Sampling," Paper presented at 1987 U. S. EPA/APCA Symposium: Measurement of Toxic and Related Air
Pollutants, Research Triangle Park, NC, 3-6 May 1987.
7. Braman, R. S., et al., "Tungstic Acid for Preconcentration and Determination of Gaseous and Particulate
Ammonia and Nitric Acid in Ambient Air," Analyt. Chem., Vol. 54:358-364, 1983.
8. | Ferm, M., Concentration Measurements and Equilibrium Studies of Ammonium, Nitrate and Sulphur
Species in Air and Precipitation, Doctoral Thesis, Department of Inorganic Chemistry, Goteborg University,
Goteborg, Sweden, 1986.
9. Ferm, M., and Sjodin A., "A Sodium Carbonate Coated Denuder for Determination of Nitrous Acid in
the Atmosphere," Atmos. Environ., Vol. 19:979-985, 1985.
10. Riggin, R. M., Technical Assistance Document for Sampling and Analysis of Toxic Organic Compounds
in Ambient Air, U.S. Environmental Protection Agency, EPA Office of Research and Development,
Atmospheric Research and Exposure Assessment Laboratory, Research Triangle Park, NC, 27511,
EPA-600/4-83-027, November, 1983.
11. | Stevens, R. K., and Rickman, E., Jr., "Research Protocol/Method for Ambient Air Sampling with
Annular Denuder Systems," U.S. Environmental Protection Agency, EPA Office of Research and
Development, Atmospheric Research and Exposure Assessment Laboratory, Research Triangle Park, NC,
27511, ASRL-ACPD-RPM 003, January 1988.
12. 40 CFR Part 58, Appendix A, B.
13. Quality Assurance Handbook for Air Pollution Measurement Systems, Volume II - Ambient Air Specific
Methods | (Interim | Edition), | U.S. | Environmental | Protection | Agency, | EPA | Office | of | Research | and
Development, Atmospheric Research and Exposure Assessment Laboratory, Research Triangle Park, NC,
27511, EPA-600/R-94/038b, May, 1972.
14. Febo H., Perino, C., and Cortiello, M., "A Denuder Technique for the Measurement of Nitrous Acid
in Urban Atmospheres," Atmos. Environ., Vol. 27:1721-1728, 1993.
15. Winberry, W.T., Jr., "Determination of the Strong Acidity of Atmospheric Fine Particles (
Using Annular Denuder Technology," U.S. Environmental Protection Agency, EPA Office of Research and
Development, Atmospheric Research and Exposure Assessment Laboratory, Research Triangle Park, NC,
27511, EPA-600/R-93-D37, November 1992.
Page 4.2-38 |  | Compendium of Methods for Inorganic Air Pollutants | [7 empty] | June 1999

## Page 43
Name: “Table 1 of 1 on page 43 of 68”
Table: 43
Chapter IO-4 |  |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
TABLE 1. ESTIMATED DETECTION AND QUANTIFICATION LIMITS
FOR THE ANNULAR DENUDER SYSTEM1
Sampling period
Detection limits, μg/m3 | 1 hour | 1 day | 1 week
Gaseous species
SO2 | 3.1 | 0.13 | 0.02
HNO3 | 2.0 | 0.08 | 0.01
NH 3 | 5.6 | 0.25 | 0.04
Particulate matter
SO4= | 1.6 | 0.07 | 0.01
NO3- | 1.8 | 0.08 | 0.01
H + | 1.0 | 0.05 | 0.001
Quantification limits, Fg/m3
Gaseous species
SO2 | 10.4 | 0.43 | 0.06
HNO3 | 6.8 | 0.28 | 0.04
NH 3 | 20.0 | 0.83 | 0.12
Particulate matter
SO4 = | 5.3 | 0.22 | 0.03
NO3- | 6.1 | 0.25 | 0.04
H + | 2.1 | 0.15 | 0.01
1Samples analyzed by ion chromatography. | Detection limits are taken as three standards deviations above field
blanks. | Quantification limits are taken as ten standard deviations above field blanks. |  | Both the detection | and
quantification limits were estimated assuming that the variance is independent of concentration.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  | Page 4.2-39

## Page 44
Name: “Table 1 of 1 on page 44 of 68”
Table: 44
Method IO-4.2 |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  | Atmospheric Acidic
TABLE 2. ACCELERATOR JET DIAMETERS AND CORRESPONDING REYNOLDS
NUMBER (RE) FOR SELECTED FLOW RATES TO OBTAIN 2.5 μm DIAMETER
(AERODYNAMIC) PARTICLES SEPARATION
Flow rate,
| L/min | Jet diameter, mm | Reynolds number (RE)
| 1.0 | 1.55 | 900
| 2.0 | 1.97 | 1,400
| 5.0 | 2.65 | 2,700
| 10.0 | 3.33 | 4,200
| 12.0 | 3.55 | 4,700
| 15.0 | 3.85 | 5,500
| 16.7 | 4.00 | 6,000
| 20.0 | 4.25 | 6,600
Page 4.2-40 | Compendium of Methods for Inorganic Air Pollutants |  | June 1999

## Page 45
Name: “Table 1 of 1 on page 45 of 68”
Table: 45
Chapter IO-4 |  |  |  | Method IO-4.2
Atmospheric Acidic |  |  | Acidic/Basic Constituents
TABLE 3. SUMMARY OF KEY PROBE SITING CRITERIA FOR ACID AEROSOL
MONITORING STATIONS
Factor |  |  | Criteria
Vertical spacing above ground | • | Representative of the breathing zone and avoiding effects of
|  | obstruction, obstacles, and roadway traffic. |  | Height of probe
intake above ground in general, 2-3 m above ground and 2-15 m
above ground in the case of nearby roadways.
| • | About 1 m or more above the structure where the sampler is
located.
Horizontal spacing from | • | Minimum horizontal separation from obstructions such as trees
obstruction and obstacles |  | should be > 20 m from the dripline and must be 10 m from the
dripline when the trees act as an obstruction.
| • | Distance from sampler inlet to an obstacle such as a building must
be at least twice the height the obstacle protrudes above the
sampler.
| • | If a sampler is located on a roof or other structures, there must be
a minimum of 2 m separation from walls, parapets, penthouses,
etc.
| • | There must be sufficient separation between the sampler and a
|  | furnace or incinerator flue. | The separation distance depends on
the height and the nature of the emissions involved.
Unrestricted airflow | • | Unrestricted airflow must exist in an arc of at least 270 degrees
around the sampler, and the predominant wind direction for the
monitoring period must be included in the 270 degree arc.
| • | A sufficient separation must exist between the sampler and nearby
Spacing from roads |  | roadways to avoid the effect of dust re-entrainment and vehicular
emissions on the measured air concentrations.
| • | Sampler should be placed at a distance of 5-25 m from the edge of
the nearest traffic lane on the roadway depending on the vertical
placement of the sampler inlet which could be 2-15 m above
ground.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  |  | Page 4.2-41

## Page 46
Name: “Table 1 of 1 on page 46 of 68”
Table: 46
Method IO-4.2 |  |  |  | Chapter IO-4
Acidic/Basic Constituents |  |  |  | Atmospheric Acidic
TABLE 4. DILUTION RATES
|  | Volume of | Equivalent
|  | 1.0 N H2SO 4 | Strong Acid
| Standard H 2SO 4 | Added to Each | Mass Collected
| Flask No. | Flask, mL | on Filter, Fg a | Approximate pH
| 1 | 0.000 | 0 | 4.09
| 2 | 0.025 | 4.90 | 4.01
| 3 | 0.050 | 9.80 | 3.95
| 4 | 0.100 | 19.60 | 3.84
| 5 | 0.200 | 39.20 | 3.68
| 6 | 0.400 | 78.40 | 3.48
| 7 | 0.800 | 156.80 | 3.23
aBased on 6.2 mL extraction volume.
TABLE 5. OTHER CHEMICAL SPECIES
AMENDABLE TO ANNULAR DENUDER TECHNOLOGY
Chemical Species
Collected on Annular
Symbol | Denuder |  | Coating | Comments
O 3 | Ozone |  | 1% (wt/wt) | Analysis by IC or
|  | NaHCO3/glycerin |  | colorimetry
HCl | Hydrogen Chloride | Na2CO3-alkali coating |  | Extracted with water/
HF | Hydrogen Fluoride |  |  | analyzed by IC
PAHs | Semi-volatile polycyclic |  | Ground | Soxhlet
SVOCs | aromatic hydrocarbons | XAD-4/glycerin |  | extraction/GC/MS
analysis
Page 4.2-42 | Compendium of Methods for Inorganic Air Pollutants |  |  | June 1999

## Page 47
Name: “Table 1 of 1 on page 47 of 68”
Table: 47
Chapter IO-4 | Method IO-4.2
Atmospheric Acidic | Acidic/Basic Constituents
es
Figure 1. Schematic view of Annular Denuder showing species collected.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-43

## Page 48
Name: “Table 1 of 1 on page 48 of 68”
Table: 48
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
Figure 2. Annular Denuder system in field sampling box.
Page 4.2-44 | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 49
Name: “Table 1 of 1 on page 49 of 68”
Table: 49
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic | Acidic/Basic Constituents
Figure 3. Available elutriator and acceleration jet assemblies.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-45

## Page 50
Name: “Table 1 of 1 on page 50 of 68”
Table: 50
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents | Atmospheric Acidic
Figure 4. Glass Annular Denuder with inset impactor assembly.
Page 4.2-46 | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 51
Name: “Table 1 of 1 on page 51 of 68”
Table: 51
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
Figure 5. Side view impactor/coupler assembly.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-47

## Page 52
Name: “Table 1 of 1 on page 52 of 68”
Table: 52
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
Figure 6. Schematic view of Annular Denuder with cyclone
adaptor for removal of coarse particles.
Page 4.2-48 | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 53
Name: “Table 1 of 1 on page 53 of 68”
Table: 53
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
Figure 7. Internal schematic of Annular Denuder.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-49

## Page 54
Name: “Table 1 of 1 on page 54 of 68”
Table: 54
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
Figure 8. Filter pack assembly.
Page 4.2-50 | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 55
Name: “Table 1 of 1 on page 55 of 68”
Table: 55
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
Figure 9. Drying train and manifold.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-51

## Page 56
Name: “Table 1 of 1 on page 56 of 68”
Table: 56
Method IO-4.2 | Chapter IO-4
Acidic/Basic Constituents | Atmospheric Acidic
Figure 10. Annular Denuder system with cyclone in heated sampling case.
Page 4.2-52 | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 57
Name: “Table 1 of 1 on page 57 of 68”
Table: 57
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic |  | Acidic/Basic Constituents
Figure 11. Frit removal from pin.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-53

## Page 58
Name: “Table 1 of 1 on page 58 of 68”
Table: 58
Method IO-4.2 | Chapter IO-4
Acidic/Basic Constituents | Atmospheric Acidic
Figure 12. Side view impactor/coupler assembly with disc removal tools.
Page 4.2-54 | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 59
Name: “Table 1 of 1 on page 59 of 68”
Table: 59
Chapter IO-4 |  |  |  |  | Method IO-4.2
Atmospheric Acidic |  |  |  | Acidic/Basic Constituents
GENERAL
Project: |  |  | Date:
Site: |  |  | Location of Sampler:
Location:
Operator:
Sample Code:
| EQUIPMENT |  |  | Sampler
Mass Flow
Controller No.: |  |  | Sodium Carbonate Denuder No.
Lab Calibration Date: |  |  | Citric Acid Denuder No.:
Flow Rate Set Point: |  |  | Filter Assembly No.:
Calibrated By:
Rotameter No.:
DGM No.:
| SAMPLING DATA |  |  | Stop
Time
Time:
Flow Rate:
Temperature:
Pressure:
Avg. Flow Rate:
Leak Check (Before):
(After):
Total Sample Vol.:
Flow Maintained Rate: |  | (±5%)
| Flow | Ambient | Barometric | Relative
| Rate (Q), | Temperature, | Pressure, | Humidity, %
Time | L/min | EC | mm Hg |  | Comments
Figure 13. Annular Denuder Field Test Data Sheet.
June 1999 | Compendium of Methods for Inorganic Air Pollutants |  |  |  | Page 4.2-55

## Page 60
Name: “Table 1 of 1 on page 60 of 68”
Table: 60
Method IO-4.2 | Chapter IO-4
Acidic/Basic Constituents | Atmospheric Acidic
Analytical
Figure 14. Major components of a commercially available ion chromatographer.
Page 4.2-56 | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 61
Name: “Table 1 of 1 on page 61 of 68”
Table: 61
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic | Acidic/Basic Constituents
Figure 15. Chromatograms of denuder/filter extract performed by the ion chromatography.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-57

## Page 62
Name: “Table 1 of 1 on page 62 of 68”
Table: 62
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents | Atmospheric Acidic
Figure 15 (Cont'd). Chromatograms of denuder/filter extract performed by the ion chromatography.
Page 4.2-58 | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 63
Name: “Table 1 of 1 on page 63 of 68”
Table: 63
Chapter IO-4 | Method IO-4.2
Atmospheric Acidic | Acidic/Basic Constituents
Figure 16. Technicon autoanalyzer flow diagram for ammonia analysis.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-59

## Page 64
Name: “Table 1 of 1 on page 64 of 68”
Table: 64
Method IO-4.2 | [7 empty] | Chapter IO-4
Acidic/Basic Constituents | [7 empty] | Atmospheric Acidic
Determination of the Strong Acidity of Atmospheric Fine-Particles (
| Name: |  |  |  | Sample I.D.:
| Date: |  |  |  | Location:
LAB:
RUN NUMBER
| Constituent | 1 | 2 | 3 | 4 | 5 | 6 | 7
pH 7 Buffer
1
2
3
pH 4 Buffer
1
2
3
EA Solution
1
2
3
Working Standards
1A1
1A2
EA
1B1
1B2
EA
2A1
2A2
Temp.
EA
2B1
2B2
EA
3A1
3A2
EA
Temp.
3B1
3B2
EA
Figure 17. pH Analytical Laboratory Log Form
Page 4.2-60 |  | Compendium of Methods for Inorganic Air Pollutants | [5 empty] | June 1999

## Page 65
Name: “Table 1 of 1 on page 65 of 68”
Table: 65
Chapter IO-4 | [6 empty] | Method IO-4.2
Atmospheric Acidic | [5 empty] | Acidic/Basic Constituents
RUN NUMBER
Constituent | 1 | 2 | 3 | 4 | 5 | 6 | 7
4A1
4A2
EA
4B1
4B2
Temp.
EA
5A1
5A2
EA
5B1
5B2
EA
6A1
6A2
Temp.
EA
6B1
6B2
EA
EA
7A1
7A2
EA
7B1
7B2
Temp.
Sample Extracts
A
A1
B
B1
C
C1
D
D1
E
E1
EA
Temp.
EA Solution
1
2
3
pH4 Buffer
1
2
3
Figure 17 (cont). pH Analytical Laboratory Log Form
June 1999 | Compendium of Methods for Inorganic Air Pollutants | [5 empty] | Page 4.2-61

## Page 66
Name: “Table 1 of 1 on page 66 of 68”
Table: 66
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
Figure 18. Field Particle personal sampler.
Page 4.2-62 | Compendium of Methods for Inorganic Air Pollutants | June 1999

## Page 67
Name: “Table 1 of 1 on page 67 of 68”
Table: 67
Chapter IO-4 |  | Method IO-4.2
Atmospheric Acidic | Acidic/Basic Constituents
Figure 19. Annular Denuder system with parallel fine particle sampler.
June 1999 | Compendium of Methods for Inorganic Air Pollutants | Page 4.2-63

## Page 68
Name: “Table 1 of 1 on page 68 of 68”
Table: 68
Method IO-4.2 |  | Chapter IO-4
Acidic/Basic Constituents |  | Atmospheric Acidic
This page intentionally left blank
Page 4.2-64 | Compendium of Methods for Inorganic Air Pollutants | June 1999